## Slide 1
Recent Advances in Through Vial Impedance Spectroscopy (TVIS) for Process Parameter Determination
Geoff Smith
Leicester School of Pharmacy, De Montfort University, United Kingdom
LyoTalk Dublin, May 22 2017
Notes: Edited in version 14

## Slide 2: Outline
Through Vial Impedance Spectroscopy (TVIS)
  Description of Measurement System
TVIS Application in Freezing
  Ice Formation and phase separation
TVIS Application in Annealing
  Surrogate Temperature Calibration
TVIS Application in Primary Drying
  Drying Rate
  Product Resistance (RP )
  Micro-collapse
Acknowledgements
Notes: Edited in version 14

## Slide 3
Through Vial Impedance Spectroscopy (TVIS)
Description of Measurement System

## Slide 4: Introduction to the TVIS System
Impedance measurements across a vial rather than within the vial
Hence “Through Vial Impedance Spectroscopy”
Features
  Single vial “non-product invasive”
  Both freezing and drying characterised in a single technique
  Non-perturbing to the packing of vials
  Stopper mechanism unaffected
TVIS
SV product temperature | 
SV sublimation rate | 
SV end point (At-Ap!) | 
Notes: V4 slide modified

## Slide 5
Freeze drying chamber
Stimulating voltage
Resultant current
LyoViewTM analysis software
LyoDEATM measurement software
Junction
box
TVIS system
(I to V convertor)
Pass through
TVIS
measurement vial
Notes: Version 14 change : don’t animate first picture – make it appear with the slide
Change photo for the pass-through to the new one
V4 edited slide

## Slide 6
Through Vial Impedance Spectroscopy (TVIS)
Theory

## Slide 7: Interfacial Polarization Phenomena
Interfacial or space charge polarization is one type of dielectric polarization.
It refers to the accumulation of charges at an interface between two dielectric materials or between two regions within a material when an external field applied.
This phenomenon occurs when an electric field is applied to a glass vial (a dielectric material) containing a liquid and/or solid (a dielectric material with some conductivity).
-
-
-
-
+
+
+
+
+
+
-
-
+
-
+
+
+
+
+
+
-
-
-
-
-
-
Accumulation of charge
Electric field
Electrodes
Absence of Electric Field
Applied Electric Field
Notes: Modified in version 4

## Slide 8: Equivalent electrical circuit model
An equivalent electrical circuit model is created by combining the circuit elements which includes the solution resistance () and the the capacitances of the glass-solution interface () and the solution () in an appropriate configuration of series and parallel elements.
CG is the capacitance of the glass-solution interface,
CS and RS are the capacitance and resistance of the solution
Notes: Version 14 I have changed CG definition from the glass wall capacitance to the interfacial capacitance

## Slide 9: Dielectric loss spectrum
As the frequency increase, increases to maximum ()
a frequency of
If > then
Which explains the sensitivity of to the height of the ice layer
Chart: Log Frequency

## Slide 10
TVIS Applications Freezing, Annealing, Primary Drying

## Slide 11: TVIS Response Surface
Imaginary Part of Capacitance
Real Part of Capacitance
High frequency
Liquid state
Liquid state
Frozen solid
Frozen solid
low frequency
Annealing = Re-heating and Re-cooling
Re-heating
Re-heating
Intermediate frequency
Re-cooling
Re-cooling
low frequency
Primary drying
Primary drying
low frequency
Notes: Full Shelf Water Study_010716_Ch1 (Filling volume = 2 mL) touch base electrode

## Slide 12: Through Vial Impedance Spectroscopy (TVIS)
The capacitance spectrum depends on both the electrical resistance and electrical capacitance of the vial contents.
Data viewing software (LyoView ™) identifies the peak frequency (FPEAK) and peak amplitude (CʺPEAK) in the imaginary part of the capacitance spectrum from which various physical properties can be determined
TVIS parameter | Application | Notes (requirements/assumptions)
FPEAK | temperature & phase (ice & eutectic formation, phase separation) | FPEAK temperature calibration (annealing stage required)
d CʺPEAK /dt | drying rate surrogate | 80% of 1o drying (assumes flat ice front)
C’ (~ 100 kHz) | end point of 1o drying | C’ (real part of the complex capacitance) is highly sensitive to low ice volumes
Notes:   Version 14 : I have moved the position of this slide and edited the colour of the last row as I don’t intend to discuss this with any data
  Slide modified in version 4
  FPEAK can be used to monitor phase behaviour (ice formation, glass transitions) and product temperature
  CʺPEAK can be used to monitor the amount of ice remaining during primary drying, from which the drying rate and end point may be determined
  Additionally. It may be used to investigate the collapse behaviour of product
  C’ at high frequency can be also used to determine the end point of primary drying

## Slide 13
TVIS Application Freezing Step
5%w/v Lactose in deionised water

## Slide 14: The relationship between FPEAK and Product Temperature/Phase Behaviour
Both resistance and capacitance parameters impact the peak frequency
However the sample resistance () has a greater temperature coefficient so the peak frequency is especially sensitive to the electrical resistance of the product.
It follows that FPEAK can be used to monitor (i) the phase behaviour and (ii) the temperatures of both the liquid and solid states.
During the solidification process the increased resistance of the frozen phase shifts the peak frequency (FPEAK) by two orders of magnitude.
During the annealing of the frozen phase, a temperature ramp of 40 oC can shift the peak frequency (FPEAK) by one order of magnitude.
Notes: Edited in version 14

## Slide 15: Phase separation
0
1
2
3
4
Time / h
6
4
3
2
1
Log Frequency
5
0.0
0.2
0.4
0.6
0.8
dielectric loss
-C" /pF
Chart: 
| 6 | 5.9030899869919438 | 5.8239109124112955 | 5.7569630370485525 | 5.6020599913279625 | 5.5228744023138026 | 5.4259741609193757 | 5.3010299956639839 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521985 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639839 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.9004927554293713 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463871 | 3.4001233554370822 | 3.300160536951354 | 3.2001388538573829 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279816 | 2.7000023971422542 | 2.6000249038929879 | 2.5000030680516945 | 2.4000023459279576 | 2.3000082025538129 | 2.2000018654066036 | 2.0999912335446838 | 2 | 1.9000009651534702 | 1.8000018278179688 | 1.7000023971422538 | 1.600003086627741 | 1.5000030680516931 | 1.4000023459279558 | 1.3000082025538131 | 1.200001865406602 | 1.0999912335446831 | 1
Time 94 min | 0.013731 | 0.017927 | 0.020119 | 0.02302 | 0.033071 | 0.039598 | 0.050917 | 0.070701 | 0.092401 | 0.11839 | 0.15752 | 0.20341 | 0.26427 | 0.34057 | 0.43852 | 0.55132 | 0.67314 | 0.78745 | 0.87791 | 0.92758 | 0.93169 | 0.89759 | 0.83915 | 0.77011 | 0.70192 | 0.6391 | 0.58187 | 0.52892 | 0.47848 | 0.42992 | 0.38264 | 0.3375 | 0.29582 | 0.25867 | 0.22659 | 0.20266 | 0.1817 | 0.16767 | 0.15365 | 0.1459 | 0.14078 | 0.13727 | 0.13846 | 0.14146 | 0.14371 | 0.14549 | 0.14911 | 0.15071 | 0.15123 | 0.14924 | 0.14113
Time 98 min | 0.0064201 | 0.0086613 | 0.0086124 | 0.0097459 | 0.013845 | 0.016144 | 0.020705 | 0.029154 | 0.038385 | 0.049362 | 0.065668 | 0.084376 | 0.10927 | 0.14094 | 0.18433 | 0.2403 | 0.3128 | 0.40067 | 0.50202 | 0.60657 | 0.70146 | 0.76925 | 0.80275 | 0.80219 | 0.77764 | 0.74174 | 0.70461 | 0.67277 | 0.64736 | 0.62503 | 0.6022 | 0.57461 | 0.53879 | 0.49613 | 0.4482 | 0.40006 | 0.35328 | 0.31062 | 0.27124 | 0.239 | 0.21314 | 0.19167 | 0.17672 | 0.16856 | 0.1609 | 0.15727 | 0.15731 | 0.15911 | 0.16239 | 0.16613 | 0.16599
Time 100 min | 0.0059865 | 0.0075857 | 0.0072824 | 0.008105 | 0.012244 | 0.012873 | 0.016318 | 0.022795 | 0.029876 | 0.038294 | 0.050775 | 0.064986 | 0.08379 | 0.1077 | 0.14023 | 0.18235 | 0.23772 | 0.30732 | 0.39245 | 0.48769 | 0.58477 | 0.66746 | 0.72423 | 0.74867 | 0.74397 | 0.72053 | 0.68991 | 0.66175 | 0.64091 | 0.6281 | 0.61921 | 0.61072 | 0.59641 | 0.57223 | 0.53804 | 0.49629 | 0.44863 | 0.39848 | 0.35278 | 0.30732 | 0.26979 | 0.23693 | 0.2144 | 0.19811 | 0.18257 | 0.17318 | 0.16853 | 0.16702 | 0.16823 | 0.17088 | 0.17195
Time 102 min | 0.0041841 | 0.0061672 | 0.0055669 | 0.0062244 | 0.009775 | 0.0098776 | 0.012622 | 0.017794 | 0.023448 | 0.030052 | 0.039761 | 0.050669 | 0.064944 | 0.082912 | 0.10719 | 0.13853 | 0.18001 | 0.23323 | 0.30082 | 0.38136 | 0.47112 | 0.55747 | 0.62885 | 0.67435 | 0.68951 | 0.68024 | 0.6563 | 0.62915 | 0.60669 | 0.59378 | 0.59078 | 0.59403 | 0.59985 | 0.60167 | 0.59462 | 0.57622 | 0.54642 | 0.5038 | 0.45278 | 0.40438 | 0.35557 | 0.31399 | 0.2761 | 0.24742 | 0.22452 | 0.20582 | 0.19445 | 0.1877 | 0.18427 | 0.18432 | 0.18342
Time 108 min | 0.0032682 | 0.0050835 | 0.0041421 | 0.0045823 | 0.0073521 | 0.0070345 | 0.0091399 | 0.013167 | 0.017603 | 0.022744 | 0.030272 | 0.038648 | 0.049479 | 0.062926 | 0.08089 | 0.10383 | 0.13406 | 0.17309 | 0.22386 | 0.28696 | 0.36247 | 0.44303 | 0.52 | 0.58189 | 0.61847 | 0.62852 | 0.61667 | 0.59235 | 0.56564 | 0.54448 | 0.53231 | 0.53058 | 0.53866 | 0.55292 | 0.56942 | 0.58375 | 0.59007 | 0.58445 | 0.56231 | 0.52727 | 0.48409 | 0.43455 | 0.38713 | 0.34172 | 0.30446 | 0.27021 | 0.24515 | 0.22664 | 0.21292 | 0.20461 | 0.19861
Log Frequency
2
3
4
5
6
1
Higher Temperature
Lower Temperature
Nucln + 24 min
Nucln + 30 min
Peak 2 shoulder appears
Nucln + 32 min
Peak 2 visible
Liquid state
Frozen solid
Freezing Step
Nucln + 38min Peak 2 shifts
to low frequency
Notes: Version 14 – I have edited the title on this slide, because the data to the right mainly concerns the latter part of the freezing process when there is sufficient ice to form thin layers of unfrozen fraction – that is responsible for the appearance of the second peak.
Version 11 - I assume that your -11.1 C (light blue line) is in fact -1.1 C (I checked process data log, it showed -11.1C)
Given that these spectra are acquired during the process of ice formation then I don’t think we can quote the TC temperatures. And we also need to register that
Full shelf 5%w/v Lactose in double distilled water 07Dec16-CH4 (Fill factor 0.7 = 3.1 g)-3mm above base electrode
Time /min Time/hr. TC2
94 1.57 -1.1
96 1.60 -1.6
98 1.63 -2.3
100 1.67 -3.2
102 1.70 -4
104 1.73 -5.1
106 1.77 -5.6
108 1.80 -6.3

## Slide 16: Ice and the unfrozen fraction
RICE
CICE
CUNFROZEN
RUNFROZEN
CG
Glass wall & solution interface
Ice
crystal
Unfrozen
fraction
The glass interface has very high resistance;
therefore it behaves primarily as capacitor (CI)
Microstructure
Ice has a conductivity (RICE) and dielectric constant (CICE) due to the percolation of protons (H+)
A layers of unfrozen fraction between the ice crystals have dielectric constants and a conductivities which are reflected in CUNFROZEN and RUNFROZEN, respectively.
The latter is strongly dependent on mobile charge carriers and hence very sensitive to the viscosity (temperature and water content) of the unfrozen fraction .
CINTERFACE

## Slide 17: Ice Formation
The thermocouple vial nucleates later than the TVIS vial
C"PEAK may provide a more reliable end point for solidification
Chart: Time / h
Chart: Log FPEAK (ice)
Chart: Temperature / oC
Ice nucleation (TVIS) @ 70 min
Solidification
end point (TVIS)
phase change only
phase change + temperature change
phase changes + temperature change
Peak
(ice)
Unfrozen fraction
Unfrozen fraction peak separates from ice peak
Notes: Edited in version 11
Full shelf 5%w/v Lactose in double distilled water 07Dec16-CH4 (Fill factor 0.7 = 3.1 g)-3mm above base electrode
TC vial nucleates after tvis (cpeak shows nucleation more clear than fpeak) presumably because TC provides heat to the liquid keeping it warmer and delaying the onset of nucleation. The reduction in the trajectory of Fpeak with time (or is it a transient increase in fpeak? – difficult to see on this scale) is what coincides with Cpeak but the onset is much less dramatic than Cpeak.
The start and end points of the freezing stage (via the spike in C"peak - in which the start and end points are quite clear - as opposed to the use of TC and F"peak data - which works for the start but is poor at determining the end point).
C"peak is sensitive to the structural changes that happens when ice forms (I think we are seeing a pecolation type phenomenon during the process of ice formation) but is much less sensitive to temperature - whereas Fpeak and TC are both sensitive to temperature and so its less clear when the end point is (as the temperature of the product is decreasing both from the decrease in shelf temperature and from the decrease on product temperature - because no more ice is forming (and the exo-thermic process has finished).
Also Fpeak has two steps – could the start of step 2 (at 1.666 h = 100 min) be when the second peak emerges – which s the point at which the unfrozen fraction viscosity starts to increase dramatically

## Slide 18: Glass Transition (Tgʹ) Determination
Chart: 1000/T (K)
| 4.4296788482835003 | 4.4120891242003104 | 4.3907793633369865 | 4.3734966105401334 | 4.3506634761801184 | 4.3318171973142814 | 4.3094160741219545 | 4.2872454448017194 | 4.2653017701002351 | 4.2417815482502652 | 4.220299641274539 | 4.1990342221288985 | 4.1744938426215805 | 4.1536863966770445 | 4.1296716911005582 | 4.1093075816724891 | 4.0858018386108279 | 4.0625634775543373 | 4.0395879620278725 | 4.0201005025125625 | 3.9992001599680025 | 3.9785160135269551 | 3.9596119580281135 | 3.9408866995073897 | 3.9192631785224381 | 3.9070130884938465 | 3.8933229511387974 | 3.8842493688094781 | 3.8767202946307391 | 3.8707180181923788 | 3.8647342995169121
Unfrozen fraction | 22.0313 | 21.981 | 21.9288 | 21.8645 | 21.806 | 21.751 | 21.6777 | 21.5918 | 21.5013 | 21.4036 | 21.2996 | 21.189 | 21.0727 | 20.9525 | 20.8199 | 20.6769 | 20.517 | 20.3563 | 20.1867 | 20.0058 | 19.804 | 19.583 | 19.3351 | 19.0636 | 18.7557 | 18.4618 | 18.2039 | 17.9839 | 17.8034 | 17.6461 | 17.5186
ln R1B | 22.0313 | 21.981 | 21.9288 | 21.8645 | 21.806 | 21.751 | 21.6777 | 21.5918 | 21.5013 | 21.4036 | 21.2996 | 21.189 | 21.0727 | 20.9525 | 20.8199 | 20.6769 | 20.517 | 20.3563 | 20.1867 | 20.0058 | 19.804 | 19.583 | 19.3351 | 19.0636 | 18.7557 | 18.4618 | 18.2039 | 17.9839 | 17.8034 | 17.6461 | 17.5186
ln R1B | 22.0313 | 21.981 | 21.9288 | 21.8645 | 21.806 | 21.751 | 21.6777 | 21.5918 | 21.5013 | 21.4036 | 21.2996 | 21.189 | 21.0727 | 20.9525 | 20.8199 | 20.6769 | 20.517
R1A
C1A
R1B
C1B
C2
Linear scale
Log scale
5.83 h
6.33 h
Initial (5.33 h)
5% w/v Lactose solution
Below Tgʹ the changes in product resistance follows Arrhenius
Above Tgʹ VTF function models the resistance profile.
-9.99 -16.74 -23.15 -29.25 -35.05 -40.59 -45.88 -50.93 oC
Notes: Full shelf 5%w/v Lactose in ultrapure water 25Nov16_CH4

## Slide 19: Glass Transition (Tgʹ) Determination
Chart: 1000/T (K)
| 4.4296788482835003 | 4.4120891242003095 | 4.3907793633369927 | 4.3734966105401272 | 4.3506634761801184 | 4.3318171973142743 | 4.3094160741219572 | 4.287245444801715 | 4.2653017701002351 | 4.2417815482502652 | 4.2202996412745311 | 4.199034222128911 | 4.1744938426215823 | 4.1536863966770516 | 4.1296716911005582 | 4.1093075816724891 | 4.0858018386108279 | 4.0625634775543373 | 4.0395879620278734 | 4.0201005025125633 | 3.9992001599680065 | 3.9785160135269551 | 3.9596119580281135 | 3.9408866995073897 | 3.9192631785224381
Unfrozen fraction | 2.8604 | 2.4482 | 3.7218 | 2.5626 | 2.9143 | 3.275 | 3.8755 | 4.121 | 4.1559 | 4.8418 | 5.1993 | 4.7401 | 5.7747 | 5.522 | 7.0254 | 6.8014 | 6.916 | 7.3811 | 9.2829 | 9.6523 | 10.685 | 13.1154 | 14.4985 | 14.238
ln R1B (Temperature above inflection point) | 7.1572 | 6.6489 | 6.8092 | 7.5837 | 8.778 | 9.9975 | 11.3906 | 12.7331 | 13.8224 | 14.6754 | 15.2593
ln R1B (Tempperature below inflection point) | 2.3457 | 2.5708 | 2.8436 | 3.0649 | 3.3572 | 3.5985 | 3.8852 | 4.169 | 4.45 | 4.7511 | 5.0261 | 5.2983 | 5.6125 | 5.8788 | 6.1863
Upper limit-Temperature above inflection point | 7.4449 | 7.2209 | 7.3355 | 7.8005 | 9.7024 | 10.0718 | 11.1045 | 13.5349 | 14.918 | 14.6575
Upper limit-Temperature below inflection point" | 3.3002 | 2.8879 | 4.1616 | 3.0023 | 3.3541 | 3.7147 | 4.3152 | 4.5607 | 4.5957 | 5.2815 | 5.639 | 5.1799 | 6.2144 | 5.9618 | 7.4651
lower limit-Temperature above inflection point" | 2.4207 | 2.0084 | 3.282 | 2.1228 | 2.4746 | 2.8352 | 3.4357 | 3.6812 | 3.7162 | 4.402 | 4.7595 | 4.3003 | 5.3349 | 5.0823 | 6.5856
Lower limit-Temperature below inflection point | 6.6059 | 6.3819 | 6.4965 | 6.9616 | 8.8634 | 9.2328 | 10.2656 | 12.6959 | 14.0791 | 13.8185
Linear scale
Log scale
5.83 h
6.33 h
Initial (5.33 h)
5% w/v Lactose solution
Below Tgʹ the changes in product resistance follows Arrhenius
Above Tgʹ VTF function models the resistance profile.
-16.74 -23.15 -29.25 -35.05 -40.59 -45.88 -50.93 oC
-28.3 C
-31.4 C
-29.8 C
Uncertainty
3.1 oC
Notes: Full shelf 5%w/v Lactose in ultrapure water 25Nov16_CH4 (analyse by limiting temperature to 19.4 c)

## Slide 20
TVIS Application Annealing for Temperature Calibration
5%w/v Lactose in deionised water

## Slide 21: The relationship between FPEAK and Product Temperature/Phase Transition
Chart: 
| 6 | 5.9030899869919438 | 5.8239109124112955 | 5.7569630370485525 | 5.6020599913279625 | 5.5228744023138026 | 5.4259741609193757 | 5.3010299956639839 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521985 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639839 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.9004927554293713 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463871 | 3.4001233554370822 | 3.300160536951354 | 3.2001388538573829 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279816 | 2.7000023971422542 | 2.6000249038929879 | 2.5000030680516945 | 2.4000023459279576 | 2.3000082025538129 | 2.2000018654066036 | 2.0999912335446838 | 2 | 1.9000009651534702 | 1.8000018278179688 | 1.7000023971422538 | 1.600003086627741 | 1.5000030680516931 | 1.4000023459279558 | 1.3000082025538131 | 1.200001865406602 | 1.0999912335446831 | 1
Time 4.77 h | 0.0019817 | 0.002341 | 0.0031922 | 0.0021042 | 0.0039425 | 0.0021922 | 0.0022352 | 0.0025687 | 0.0028841 | 0.0032143 | 0.0035773 | 0.0039388 | 0.0043811 | 0.0049232 | 0.005669 | 0.0066275 | 0.0079033 | 0.0095248 | 0.011663 | 0.014443 | 0.018072 | 0.022786 | 0.029014 | 0.037323 | 0.048245 | 0.062812 | 0.0822 | 0.10761 | 0.14035 | 0.18098 | 0.22858 | 0.28024 | 0.3316 | 0.37733 | 0.41297 | 0.43407 | 0.44075 | 0.43251 | 0.41214 | 0.38287 | 0.34804 | 0.31109 | 0.27403 | 0.24272 | 0.21066 | 0.18555 | 0.16525 | 0.14973 | 0.1389 | 0.13231 | 0.13006
Time 4.87 h | 0.0018092 | 0.0015271 | 0.0029576 | 0.0024553 | 0.0042343 | 0.0024824 | 0.0025333 | 0.0028367 | 0.0031454 | 0.0034856 | 0.0038964 | 0.0043236 | 0.0048427 | 0.0055022 | 0.0063825 | 0.0075313 | 0.009043 | 0.010962 | 0.013454 | 0.016725 | 0.021075 | 0.026717 | 0.03417 | 0.044152 | 0.057334 | 0.07493 | 0.098275 | 0.12866 | 0.16721 | 0.2137 | 0.26613 | 0.32026 | 0.37083 | 0.41183 | 0.43957 | 0.45019 | 0.4449 | 0.42573 | 0.39648 | 0.3608 | 0.32223 | 0.28516 | 0.2482 | 0.21944 | 0.1927 | 0.17275 | 0.15796 | 0.14808 | 0.14348 | 0.14364 | 0.14879
Time 4.97 h | 0.0010215 | 0.00097244 | 0.0032415 | 0.0028953 | 0.0045008 | 0.0027475 | 0.0028009 | 0.0031258 | 0.0034893 | 0.0038535 | 0.0043461 | 0.0048551 | 0.0054925 | 0.0062938 | 0.0073835 | 0.0087867 | 0.01061 | 0.01298 | 0.016107 | 0.02015 | 0.025553 | 0.032576 | 0.041927 | 0.05446 | 0.071058 | 0.093212 | 0.12242 | 0.15991 | 0.20616 | 0.25963 | 0.3167 | 0.37117 | 0.41669 | 0.44774 | 0.46073 | 0.45637 | 0.43655 | 0.40467 | 0.36739 | 0.32748 | 0.28663 | 0.25224 | 0.22069 | 0.19712 | 0.17807 | 0.16548 | 0.159 | 0.15789 | 0.16266 | 0.1733 | 0.18945
Time 5.07 h | 0.0017455 | 0.0010732 | 0.0035166 | 0.0030495 | 0.0047292 | 0.0029943 | 0.0030815 | 0.0034916 | 0.0039372 | 0.0043701 | 0.004976 | 0.005643 | 0.0064776 | 0.0075304 | 0.0089433 | 0.010762 | 0.013165 | 0.016228 | 0.020248 | 0.025536 | 0.032619 | 0.04188 | 0.054219 | 0.070865 | 0.09294 | 0.12224 | 0.16038 | 0.2079 | 0.26374 | 0.32403 | 0.38228 | 0.4307 | 0.46269 | 0.47449 | 0.46741 | 0.44272 | 0.40889 | 0.36621 | 0.3251 | 0.28479 | 0.24949 | 0.22117 | 0.19782 | 0.18144 | 0.17546 | 0.17374 | 0.17821 | 0.18904 | 0.20679 | 0.2308 | 0.26052
Time 5.17 h | 0.0029704 | 0.0027506 | 0.0044557 | 0.0028312 | 0.0047101 | 0.0030059 | 0.0031787 | 0.0038291 | 0.0044639 | 0.005046 | 0.0058643 | 0.0067724 | 0.0079118 | 0.0093372 | 0.011259 | 0.013716 | 0.016985 | 0.02114 | 0.026682 | 0.033943 | 0.043723 | 0.056585 | 0.073823 | 0.09715 | 0.12787 | 0.16806 | 0.21858 | 0.27783 | 0.3416 | 0.40211 | 0.45056 | 0.47958 | 0.48569 | 0.47069 | 0.43961 | 0.40029 | 0.35716 | 0.31391 | 0.2752 | 0.24353 | 0.21907 | 0.20106 | 0.19033 | 0.18625 | 0.19193 | 0.20317 | 0.22211 | 0.24729 | 0.27922 | 0.31694 | 0.35889
Time 5.27 h | 0.0019891 | 0.0015314 | 0.0042182 | 0.0033614 | 0.005312 | 0.0036332 | 0.0039371 | 0.0046682 | 0.0054583 | 0.0062587 | 0.0073877 | 0.0086552 | 0.010292 | 0.012333 | 0.015058 | 0.018588 | 0.023237 | 0.029256 | 0.037254 | 0.047802 | 0.062071 | 0.080895 | 0.10617 | 0.1401 | 0.18392 | 0.23876 | 0.30247 | 0.36942 | 0.43077 | 0.4764 | 0.49931 | 0.49715 | 0.47373 | 0.43598 | 0.39007 | 0.34618 | 0.30556 | 0.27147 | 0.24515 | 0.22792 | 0.21942 | 0.22039 | 0.22814 | 0.24855 | 0.27008 | 0.30202 | 0.34096 | 0.38421 | 0.42901 | 0.47026 | 0.50189
Time 5.37 h | 0.0046317 | 0.0053022 | 0.003927 | 0.0039805 | 0.0059145 | 0.0042663 | 0.0046399 | 0.0058051 | 0.0070519 | 0.0083358 | 0.010107 | 0.012096 | 0.014595 | 0.01781 | 0.022125 | 0.027625 | 0.034866 | 0.044281 | 0.056825 | 0.073435 | 0.09596 | 0.12566 | 0.16506 | 0.21622 | 0.27832 | 0.34866 | 0.4195 | 0.47968 | 0.51895 | 0.53109 | 0.51719 | 0.48395 | 0.44098 | 0.39669 | 0.35774 | 0.32768 | 0.30853 | 0.29916 | 0.30107 | 0.31205 | 0.33471 | 0.36348 | 0.39947 | 0.44356 | 0.47454 | 0.50226 | 0.51599 | 0.51186 | 0.49017 | 0.45471 | 0.41012
Time 5.47 h | 0.0052467 | 0.0031832 | 0.0074506 | 0.0050694 | 0.0072304 | 0.0057999 | 0.006411 | 0.0080493 | 0.0098963 | 0.011931 | 0.01479 | 0.017926 | 0.022117 | 0.027348 | 0.034403 | 0.043444 | 0.055425 | 0.071116 | 0.092172 | 0.12018 | 0.15805 | 0.20693 | 0.26873 | 0.3423 | 0.42025 | 0.49276 | 0.54703 | 0.57392 | 0.57215 | 0.54717 | 0.50939 | 0.46933 | 0.43413 | 0.40896 | 0.39611 | 0.39701 | 0.40745 | 0.42893 | 0.45445 | 0.48182 | 0.50236 | 0.51493 | 0.50614 | 0.49215 | 0.45322 | 0.41004 | 0.36287 | 0.31774 | 0.27693 | 0.2416 | 0.21186
4
5
6
7
8
Time / h
6
4
3
2
1
Log Frequency
5
0.0
0.2
0.4
0.6
0.8
dielectric loss
-C" /pF
Log Frequency
2
3
4
5
6
1
Higher Temperature
Lower Temperature
-46.1 oC
-43.3 oC
-40.1oC
-21.5 oC
-36.5 oC
-29.1oC
-25.3 oC
-32.8oC
Re-heating
9
Re-heating
Notes: Full shelf 5%w/v Lactose in ultrapure water 23Jan17-CH3 (Fill factor 0.7 = 3.04 g)-3mm above base electrode

## Slide 22: The relationship between FPEAK and Product Temperature/Phase Transition
Chart: 
| 6 | 5.9030899869919438 | 5.8239109124112955 | 5.7569630370485525 | 5.6020599913279625 | 5.5228744023138026 | 5.4259741609193757 | 5.3010299956639839 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521985 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639839 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.9004927554293713 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463871 | 3.4001233554370822 | 3.300160536951354 | 3.2001388538573829 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279816 | 2.7000023971422542 | 2.6000249038929879 | 2.5000030680516945 | 2.4000023459279576 | 2.3000082025538129 | 2.2000018654066036 | 2.0999912335446838 | 2 | 1.9000009651534702 | 1.8000018278179688 | 1.7000023971422538 | 1.600003086627741 | 1.5000030680516931 | 1.4000023459279558 | 1.3000082025538131 | 1.200001865406602 | 1.0999912335446831 | 1
Time 7.7 h | 0.011193 | 0.012362 | 0.012438 | 0.013416 | 0.017154 | 0.019576 | 0.024073 | 0.032106 | 0.040694 | 0.050657 | 0.065755 | 0.083057 | 0.10619 | 0.1359 | 0.17678 | 0.23012 | 0.30035 | 0.38798 | 0.49252 | 0.60395 | 0.70881 | 0.78657 | 0.82608 | 0.825 | 0.79277 | 0.74298 | 0.68787 | 0.63498 | 0.58687 | 0.54266 | 0.4997 | 0.45632 | 0.41019 | 0.36416 | 0.32181 | 0.27905 | 0.24205 | 0.21024 | 0.18305 | 0.15975 | 0.14383 | 0.12671 | 0.11682 | 0.097155 | 0.10434 | 0.099739 | 0.097515 | 0.097491 | 0.099382 | 0.10125 | 0.10544
Time 7.8 h | 0.0086335 | 0.019002 | 0.012658 | 0.010129 | 0.012571 | 0.013961 | 0.016847 | 0.02213 | 0.02778 | 0.034154 | 0.043799 | 0.054778 | 0.069318 | 0.087802 | 0.11306 | 0.14599 | 0.19015 | 0.24776 | 0.32242 | 0.41328 | 0.51713 | 0.61939 | 0.70574 | 0.76171 | 0.78026 | 0.76689 | 0.73295 | 0.69069 | 0.64927 | 0.61264 | 0.57951 | 0.54777 | 0.51465 | 0.47689 | 0.43687 | 0.3911 | 0.34557 | 0.30226 | 0.26208 | 0.22784 | 0.19665 | 0.17204 | 0.1498 | 0.1416 | 0.12284 | 0.11338 | 0.10812 | 0.10456 | 0.10345 | 0.10292 | 0.10427
Time 7.9 h | 0.0071174 | 0.014937 | 0.011187 | 0.0076248 | 0.0099665 | 0.0095716 | 0.011342 | 0.01471 | 0.018318 | 0.022411 | 0.028246 | 0.034943 | 0.043713 | 0.054721 | 0.069611 | 0.08883 | 0.11452 | 0.14834 | 0.19367 | 0.2527 | 0.32835 | 0.417 | 0.51258 | 0.60306 | 0.67209 | 0.71086 | 0.71807 | 0.70162 | 0.67295 | 0.64161 | 0.61518 | 0.59453 | 0.57913 | 0.56449 | 0.543 | 0.51561 | 0.4807 | 0.4357 | 0.38832 | 0.34011 | 0.29281 | 0.25655 | 0.2167 | 0.19886 | 0.16762 | 0.14897 | 0.1348 | 0.12393 | 0.11736 | 0.11266 | 0.10983
Time 8 h | 0.0066875 | 0.0072971 | 0.0060739 | 0.0062197 | 0.0082996 | 0.0072917 | 0.0083212 | 0.010549 | 0.012982 | 0.015695 | 0.019476 | 0.023835 | 0.029507 | 0.036567 | 0.046034 | 0.058137 | 0.074182 | 0.095174 | 0.12333 | 0.16065 | 0.21056 | 0.2739 | 0.35132 | 0.43921 | 0.52681 | 0.60209 | 0.65358 | 0.67562 | 0.67282 | 0.65336 | 0.62868 | 0.60616 | 0.5893 | 0.57943 | 0.5733 | 0.56434 | 0.55057 | 0.52741 | 0.4929 | 0.44912 | 0.40193 | 0.35095 | 0.30786 | 0.25473 | 0.22997 | 0.20033 | 0.17498 | 0.15579 | 0.14111 | 0.13045 | 0.12252
Time 8.1 h | 0.005947 | 0.0067281 | 0.0052524 | 0.0053001 | 0.0073028 | 0.0058829 | 0.0064586 | 0.0079968 | 0.0097409 | 0.011645 | 0.014341 | 0.017255 | 0.021161 | 0.025983 | 0.032413 | 0.040564 | 0.051285 | 0.065215 | 0.083812 | 0.10848 | 0.14191 | 0.18569 | 0.24255 | 0.31373 | 0.39538 | 0.48106 | 0.55936 | 0.61864 | 0.65248 | 0.66063 | 0.65036 | 0.63026 | 0.60966 | 0.59362 | 0.58267 | 0.57619 | 0.57029 | 0.55842 | 0.54009 | 0.51084 | 0.47236 | 0.43098 | 0.38259 | 0.34715 | 0.29036 | 0.25192 | 0.21922 | 0.19172 | 0.16974 | 0.15248 | 0.13857
Time 8.2 h | 0.0054315 | 0.0061022 | 0.0046427 | 0.0046453 | 0.0065422 | 0.0049313 | 0.0052257 | 0.0063414 | 0.0076549 | 0.0090535 | 0.010998 | 0.013093 | 0.01591 | 0.019392 | 0.023988 | 0.029796 | 0.037412 | 0.047207 | 0.060222 | 0.07738 | 0.10068 | 0.13151 | 0.17261 | 0.22668 | 0.29373 | 0.37255 | 0.4567 | 0.53596 | 0.60009 | 0.6416 | 0.65935 | 0.65755 | 0.64444 | 0.62595 | 0.60857 | 0.5955 | 0.58357 | 0.57272 | 0.55827 | 0.53495 | 0.50934 | 0.46693 | 0.42882 | 0.3737 | 0.34052 | 0.29831 | 0.25978 | 0.22526 | 0.19757 | 0.1741 | 0.15593
Time 8.3 h | 0.0030829 | 0.0061846 | 0.013615 | 0.0045985 | 0.0058723 | 0.0042349 | 0.0043781 | 0.0052895 | 0.0063454 | 0.0074338 | 0.0089232 | 0.010573 | 0.01263 | 0.015242 | 0.018696 | 0.023048 | 0.028709 | 0.035973 | 0.045566 | 0.058127 | 0.075103 | 0.097557 | 0.12775 | 0.1684 | 0.22103 | 0.28704 | 0.3645 | 0.44712 | 0.52608 | 0.59127 | 0.63639 | 0.65932 | 0.66398 | 0.65564 | 0.64051 | 0.62404 | 0.6082 | 0.59178 | 0.57344 | 0.55266 | 0.52768 | 0.49131 | 0.45594 | 0.42746 | 0.37699 | 0.3334 | 0.29368 | 0.25654 | 0.22511 | 0.19735 | 0.17422
Time 8.4 h | 0.0044719 | 0.0052646 | 0.0038505 | 0.0038229 | 0.0055183 | 0.0038613 | 0.0038757 | 0.0045981 | 0.0054556 | 0.0063256 | 0.0075072 | 0.0088339 | 0.0105 | 0.012494 | 0.015202 | 0.018609 | 0.023049 | 0.028688 | 0.036059 | 0.045662 | 0.05855 | 0.075459 | 0.098272 | 0.12915 | 0.16991 | 0.22302 | 0.28916 | 0.36565 | 0.44673 | 0.52373 | 0.58797 | 0.63371 | 0.65883 | 0.66616 | 0.66058 | 0.64736 | 0.63152 | 0.61024 | 0.5896 | 0.56867 | 0.53829 | 0.51193 | 0.47677 | 0.43046 | 0.40086 | 0.36147 | 0.32353 | 0.28629 | 0.25237 | 0.22152 | 0.19517
4
5
6
7
8
Time / h
6
4
3
2
1
Log Frequency
5
0.0
0.2
0.4
0.6
0.8
dielectric loss
-C" /pF
Log Frequency
2
3
4
5
6
1
Higher Temperature
Lower Temperature
-13.9oC
-16.3 oC
-19.1oC
-35.7 oC
-22.2 oC
-28.7oC
-32.5 oC
-25.5oC
Re-cooling
9
Re-cooling
Notes: Full shelf 5%w/v Lactose in ultrapure water 23Jan17-CH3 (Fill factor 0.7 = 3.04 g)-3mm above base electrode

## Slide 23: Temperature Calibration
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124112991 | 5.7569630370485525 | 5.6020599913279625 | 5.5228744023138026 | 5.4259741609193757 | 5.3010299956639821 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521985 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639821 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.9004927554293705 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463871 | 3.4001233554370813 | 3.3001605369513531 | 3.2001388538573816 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279803 | 2.7000023971422542 | 2.600024903892987 | 2.5000030680516936 | 2.4000023459279562 | 2.3000082025538129 | 2.2000018654066023 | 2.0999912335446842 | 2 | 1.9000009651534697 | 1.8000018278179688 | 1.7000023971422538 | 1.6000030866277406 | 1.5000030680516931 | 1.4000023459279558 | 1.3000082025538131 | 1.200001865406602 | 1.099991233544684 | 1
-41.1 C | 0.0042134 | 0.0047101 | 0.0033454 | 0.0033264 | 0.0049656 | 0.0032695 | 0.0032292 | 0.0037856 | 0.0044547 | 0.0051021 | 0.0059756 | 0.0069349 | 0.0081176 | 0.0095497 | 0.01141 | 0.013797 | 0.016872 | 0.020761 | 0.02584 | 0.03238 | 0.041061 | 0.052364 | 0.067422 | 0.08781 | 0.11489 | 0.15116 | 0.19873 | 0.25829 | 0.32868 | 0.40523 | 0.48096 | 0.54761 | 0.59992 | 0.63518 | 0.65443 | 0.65802 | 0.65165 | 0.63516 | 0.61476 | 0.58962 | 0.56459 | 0.5325 | 0.50365 | 0.47897 | 0.43647 | 0.40124 | 0.36651 | 0.33155 | 0.29807 | 0.26646 | 0.23702
-37.6 C | 0.0027378 | 0.0059614 | 0.013324 | 0.0041469 | 0.0052967 | 0.0036306 | 0.0036043 | 0.0042662 | 0.0050403 | 0.0058204 | 0.0068674 | 0.0080326 | 0.0094942 | 0.011227 | 0.013598 | 0.016565 | 0.02042 | 0.025289 | 0.03167 | 0.039952 | 0.050998 | 0.065489 | 0.084915 | 0.11124 | 0.14621 | 0.19241 | 0.25123 | 0.32183 | 0.40032 | 0.47938 | 0.55041 | 0.60627 | 0.64356 | 0.6617 | 0.66466 | 0.65682 | 0.6399 | 0.62225 | 0.5996 | 0.5759 | 0.54836 | 0.51896 | 0.4881 | 0.46152 | 0.41421 | 0.37687 | 0.34023 | 0.30319 | 0.26914 | 0.23777 | 0.20928
-30.1 C | 0.0048643 | 0.0063044 | 0.013553 | 0.0048896 | 0.0063048 | 0.0046852 | 0.0049078 | 0.0059385 | 0.0071506 | 0.0084323 | 0.010207 | 0.012187 | 0.014661 | 0.017807 | 0.02197 | 0.027242 | 0.034087 | 0.042917 | 0.054615 | 0.070042 | 0.090941 | 0.11862 | 0.15567 | 0.20491 | 0.2671 | 0.34198 | 0.42489 | 0.50671 | 0.57736 | 0.62784 | 0.65526 | 0.66153 | 0.65356 | 0.63741 | 0.62106 | 0.6053 | 0.59105 | 0.57932 | 0.56279 | 0.54157 | 0.51484 | 0.48178 | 0.43923 | 0.40308 | 0.35364 | 0.31099 | 0.2716 | 0.2363 | 0.20694 | 0.18222 | 0.16178
-24.3 C | 0.0061986 | 0.0069698 | 0.0055194 | 0.0055866 | 0.0076011 | 0.0062942 | 0.007005 | 0.0087317 | 0.010672 | 0.012803 | 0.015826 | 0.019116 | 0.023505 | 0.028951 | 0.036209 | 0.045449 | 0.05762 | 0.073474 | 0.094682 | 0.12285 | 0.16089 | 0.21035 | 0.27369 | 0.35085 | 0.43614 | 0.52088 | 0.5927 | 0.64109 | 0.66233 | 0.65986 | 0.64282 | 0.62064 | 0.60022 | 0.58649 | 0.57696 | 0.57172 | 0.56683 | 0.55362 | 0.52951 | 0.49671 | 0.45429 | 0.40984 | 0.3589 | 0.31519 | 0.27283 | 0.23473 | 0.20458 | 0.17959 | 0.16011 | 0.145 | 0.13327
-20.1 C | 0.0068323 | 0.01549 | 0.01057 | 0.0071073 | 0.0093221 | 0.0086855 | 0.010157 | 0.013063 | 0.016208 | 0.019747 | 0.024777 | 0.030548 | 0.0381 | 0.047545 | 0.060274 | 0.076644 | 0.098452 | 0.12711 | 0.16558 | 0.21613 | 0.28223 | 0.36228 | 0.45334 | 0.54588 | 0.62499 | 0.67877 | 0.70149 | 0.69676 | 0.6746 | 0.64562 | 0.61937 | 0.59874 | 0.58488 | 0.5733 | 0.55968 | 0.53985 | 0.51332 | 0.47276 | 0.4288 | 0.38002 | 0.33247 | 0.28506 | 0.24881 | 0.20851 | 0.18812 | 0.16391 | 0.14694 | 0.1337 | 0.1242 | 0.11734 | 0.11336
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124112991 | 5.7569630370485525 | 5.6020599913279625 | 5.5228744023138026 | 5.4259741609193757 | 5.3010299956639821 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521985 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639821 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.9004927554293705 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463871 | 3.4001233554370813 | 3.3001605369513531 | 3.2001388538573816 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279803 | 2.7000023971422542 | 2.600024903892987 | 2.5000030680516936 | 2.4000023459279562 | 2.3000082025538129 | 2.2000018654066023 | 2.0999912335446842 | 2 | 1.9000009651534697 | 1.8000018278179688 | 1.7000023971422538 | 1.6000030866277406 | 1.5000030680516931 | 1.4000023459279558 | 1.3000082025538131 | 1.200001865406602 | 1.099991233544684 | 1
-41.1 C | 0.00073659 | 0.00041878 | 0.003422 | 0.0030545 | 0.0044768 | 0.0027076 | 0.0027368 | 0.0030339 | 0.0033684 | 0.0037141 | 0.004177 | 0.004647 | 0.005241 | 0.0059742 | 0.0069837 | 0.0082793 | 0.0099743 | 0.012173 | 0.015057 | 0.018817 | 0.02379 | 0.030272 | 0.038871 | 0.050401 | 0.065659 | 0.086016 | 0.11295 | 0.14768 | 0.19103 | 0.24198 | 0.29758 | 0.35227 | 0.40016 | 0.4354 | 0.45427 | 0.45564 | 0.44148 | 0.41348 | 0.37881 | 0.33901 | 0.3 | 0.26283 | 0.23141 | 0.20616 | 0.18231 | 0.16642 | 0.15667 | 0.15231 | 0.1537 | 0.16067 | 0.17252
-37.6 C | 0.0013146 | 0.0053088 | 0.00051647 | 0.0035613 | 0.0046302 | 0.0029154 | 0.0029801 | 0.0033607 | 0.0037764 | 0.004175 | 0.0047438 | 0.0053439 | 0.0061078 | 0.0070641 | 0.0083551 | 0.010007 | 0.012185 | 0.015002 | 0.01868 | 0.023498 | 0.029942 | 0.038347 | 0.049558 | 0.064647 | 0.084631 | 0.11125 | 0.14608 | 0.19001 | 0.24265 | 0.30099 | 0.35962 | 0.41111 | 0.44899 | 0.46835 | 0.46756 | 0.45051 | 0.41965 | 0.38137 | 0.34049 | 0.29943 | 0.26172 | 0.23176 | 0.20435 | 0.18697 | 0.17421 | 0.16864 | 0.1693 | 0.17597 | 0.18937 | 0.20859 | 0.23381
-30.3 C | 0.0031788 | 0.0037585 | 0.002653 | 0.0029709 | 0.0049413 | 0.0032758 | 0.0035487 | 0.0042989 | 0.0050358 | 0.0057549 | 0.0067704 | 0.0079011 | 0.0093443 | 0.011144 | 0.013582 | 0.016684 | 0.020783 | 0.026081 | 0.033124 | 0.042371 | 0.054872 | 0.071402 | 0.093551 | 0.12339 | 0.16235 | 0.21202 | 0.27164 | 0.33697 | 0.40067 | 0.45295 | 0.48561 | 0.49405 | 0.47992 | 0.44859 | 0.40685 | 0.36209 | 0.32026 | 0.28018 | 0.24882 | 0.22531 | 0.20895 | 0.20234 | 0.20345 | 0.21302 | 0.22675 | 0.24942 | 0.28006 | 0.31673 | 0.3589 | 0.4038 | 0.44737
-24.0 C | 0.0046998 | 0.0031445 | 0.0064846 | 0.0043939 | 0.0063456 | 0.0047445 | 0.0051374 | 0.0063975 | 0.0077973 | 0.0092851 | 0.011361 | 0.013682 | 0.016622 | 0.02038 | 0.025427 | 0.031854 | 0.040369 | 0.051443 | 0.066204 | 0.085751 | 0.11231 | 0.14724 | 0.19308 | 0.25135 | 0.31974 | 0.39338 | 0.46239 | 0.51432 | 0.54101 | 0.53945 | 0.51477 | 0.47571 | 0.43231 | 0.39187 | 0.3607 | 0.33875 | 0.33014 | 0.3312 | 0.3433 | 0.36557 | 0.39536 | 0.43074 | 0.46504 | 0.50005 | 0.5118 | 0.51349 | 0.498 | 0.46694 | 0.42565 | 0.37901 | 0.33296
-20.3 C | 0.006034 | 0.0032347 | 0.0080358 | 0.0054713 | 0.0077132 | 0.0064327 | 0.0072229 | 0.0091338 | 0.011272 | 0.013654 | 0.017047 | 0.020778 | 0.02576 | 0.032 | 0.040423 | 0.051256 | 0.065684 | 0.084631 | 0.11013 | 0.14401 | 0.18949 | 0.24742 | 0.31822 | 0.39831 | 0.47753 | 0.54398 | 0.5859 | 0.59786 | 0.5831 | 0.55078 | 0.51184 | 0.47578 | 0.44845 | 0.43286 | 0.43101 | 0.44041 | 0.4568 | 0.47976 | 0.49873 | 0.51106 | 0.51133 | 0.49523 | 0.46505 | 0.41679 | 0.38015 | 0.33339 | 0.2905 | 0.25292 | 0.22105 | 0.19498 | 0.17399
Re-cooling
Re-heating
Re-cooling
At low temperatures, the two peaks are merged, forming one peak below –35 C.
At higher temperatures > – 20 C the two curve separate to some degree
Re-heating
The two peaks on re-heating are well separated at all temperatures between – 20 C and – 50 C
Poor peak separation
Notes: Full shelf 5%w/v Lactose in ultrapure water 23Jan17-CH3 (Fill factor 0.7 = 3.04 g)-3mm above base electrode
There is good agreement between re-heating and re-cooling but the low temperature (-19 to -50) is covered better by the re-heating curve.
The two peaks in the re-cooling spectra are merged and form one peak below –30 C.
At higher temperatures > -20 C the two curve separate to some degree
In contrast, the two peaks on re-heating at quite well separated at all temperatures below -20 C

## Slide 24: Temperature Calibration
Thermal homogeneity of the frozen solid demonstrated by two TCS (top and bottom)
FPEAK profile during annealing has ‘similar’ profile with product temperature.
Assuming thermal equivalence between the TC and TVIS vial (?!!) then temperature calibration from annealing might be employed for the prediction of temperature during primary drying
Re-heating curve selected because of the wider range of temperatures for the observation of a single peak
Chart: Time / h
Chart: Log FPEAK
| 2.3709079999999996 | 2.3885049999999999 | 2.4100199999999994 | 2.4335930000000001 | 2.456945999999999 | 2.4844489999999997 | 2.5136549999999995 | 2.545086 | 2.5786149999999997 | 2.6156109999999995 | 2.6527509999999994 | 2.6938330000000001 | 2.7347070000000002 | 2.7777280000000002 | 2.8227039999999994 | 2.8670719999999998 | 2.9118399999999993 | 2.9583779999999997 | 3.0051109999999999 | 3.0520199999999997 | 3.1028879999999996 | 3.1534900000000001 | 3.203856 | 3.251633 | 3.29948
Re-heating | -47.7 | -47 | -46.1 | -45.1 | -44.2 | -43.3 | -42.1 | -41.1 | -40.1 | -38.8 | -37.6 | -36.5 | -35.4 | -34.1 | -32.8 | -31.5 | -30.3 | -29.1 | -27.8 | -26.4 | -25.3 | -24 | -22.8 | -21.5 | -20.3
Re-cooling | -16.3 | -17.1 | -18 | -19.1 | -20.1 | -21.1 | -22.2 | -23.2 | -24.3 | -25.5 | -26.6 | -27.7 | -28.7 | -30.1 | -31.2 | -32.5
Notes: Full shelf 5%w/v Lactose in ultrapure water 23Jan17-CH3 (Fill factor 0.7 = 3.04 g)-3mm above base electrode
There is good agreement between re-heating and re-cooling but the low temperature (-19 to -50) is covered better by the re-heating curve.

## Slide 25
TVIS Application Primary Drying (1) Product temperature prediction
5%w/v Lactose in deionised water

## Slide 26: Temperature Prediction in Primary Drying
Chart: Time / h
Re-heating calibration curve selected for temperature prediction in primary drying : T(FPEAK)
Good agreement between production temperature (by TC) and T(FPEAK)
At approx. ̶ 32 oC (product collapse) the time profile of each parameter undergoes some instability
Chart: Log FPEAK
| 2.3709079999999996 | 2.3885049999999999 | 2.4100199999999994 | 2.4335930000000001 | 2.456945999999999 | 2.4844489999999997 | 2.5136549999999995 | 2.545086 | 2.5786149999999997 | 2.6156109999999995 | 2.6527509999999994 | 2.6938330000000001 | 2.7347070000000002 | 2.7777280000000002 | 2.8227039999999994 | 2.8670719999999998 | 2.9118399999999993 | 2.9583779999999997 | 3.0051109999999999 | 3.0520199999999997 | 3.1028879999999996 | 3.1534900000000001 | 3.203856 | 3.251633 | 3.29948
Re-heating | -47.7 | -47 | -46.1 | -45.1 | -44.2 | -43.3 | -42.1 | -41.1 | -40.1 | -38.8 | -37.6 | -36.5 | -35.4 | -34.1 | -32.8 | -31.5 | -30.3 | -29.1 | -27.8 | -26.4 | -25.3 | -24 | -22.8 | -21.5 | -20.3
Re-cooling | -16.3 | -17.1 | -18 | -19.1 | -20.1 | -21.1 | -22.2 | -23.2 | -24.3 | -25.5 | -26.6 | -27.7 | -28.7 | -30.1 | -31.2 | -32.5
Notes: Full shelf 5%w/v Lactose in ultrapure water 23Jan17-CH3 (Fill factor 0.7 = 3.04 g)-3mm above base electrode

## Slide 27
TVIS Application Primary Drying (2) Drying Rate Prediction
5%w/v Lactose in deionised water

## Slide 28: The relationship between FPEAK and Product Temperature/Phase Transition
Chart: 
| 6 | 5.9030899869919438 | 5.8239109124112955 | 5.7569630370485525 | 5.6020599913279625 | 5.5228744023138026 | 5.4259741609193757 | 5.3010299956639839 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521985 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639839 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.9004927554293713 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463871 | 3.4001233554370822 | 3.300160536951354 | 3.2001388538573829 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279816 | 2.7000023971422542 | 2.6000249038929879 | 2.5000030680516945 | 2.4000023459279576 | 2.3000082025538129 | 2.2000018654066036 | 2.0999912335446838 | 2 | 1.9000009651534702 | 1.8000018278179688 | 1.7000023971422538 | 1.600003086627741 | 1.5000030680516931 | 1.4000023459279558 | 1.3000082025538131 | 1.200001865406602 | 1.0999912335446831 | 1
Time 11.77 h | 0.0034144 | 0.0036926 | 0.0026205 | 0.0028242 | 0.0046259 | 0.0029246 | 0.0030195 | 0.0035876 | 0.0041856 | 0.0047428 | 0.0055009 | 0.0063098 | 0.0073074 | 0.0085194 | 0.010139 | 0.01219 | 0.014841 | 0.018215 | 0.022574 | 0.028204 | 0.035664 | 0.04533 | 0.058182 | 0.075502 | 0.098591 | 0.12954 | 0.17048 | 0.22268 | 0.28585 | 0.35664 | 0.42894 | 0.49504 | 0.54912 | 0.58706 | 0.60994 | 0.61796 | 0.61422 | 0.60221 | 0.58855 | 0.56696 | 0.54048 | 0.52001 | 0.49065 | 0.46085 | 0.43641 | 0.40685 | 0.38162 | 0.34516 | 0.31711 | 0.28926 | 0.26056
Time 14.77 h | 0.0034331 | 0.0034293 | 0.0044876 | 0.0029176 | 0.00446 | 0.0028144 | 0.002755 | 0.003176 | 0.0037811 | 0.0043915 | 0.0052537 | 0.0062128 | 0.0074055 | 0.008856 | 0.010794 | 0.013219 | 0.016396 | 0.020453 | 0.025798 | 0.032801 | 0.042223 | 0.054646 | 0.071435 | 0.094338 | 0.12493 | 0.16564 | 0.21792 | 0.28085 | 0.35066 | 0.42017 | 0.48102 | 0.52625 | 0.55311 | 0.56263 | 0.55807 | 0.54651 | 0.52901 | 0.50877 | 0.49017 | 0.47388 | 0.4523 | 0.43434 | 0.41361 | 0.38145 | 0.36385 | 0.33718 | 0.3133 | 0.28819 | 0.2653 | 0.2455 | 0.22764
Time 16.77 h | 0.0038166 | 0.0042533 | 0.0041977 | 0.0029457 | 0.0044631 | 0.0029017 | 0.0028807 | 0.003246 | 0.0038638 | 0.0045325 | 0.0054656 | 0.0065208 | 0.0078649 | 0.009537 | 0.011767 | 0.014609 | 0.018396 | 0.023343 | 0.030011 | 0.038902 | 0.051143 | 0.067551 | 0.089946 | 0.12051 | 0.16074 | 0.21192 | 0.27251 | 0.33728 | 0.39849 | 0.44794 | 0.48022 | 0.49447 | 0.49416 | 0.48339 | 0.46919 | 0.45466 | 0.4425 | 0.43018 | 0.41917 | 0.40742 | 0.39201 | 0.37233 | 0.34936 | 0.33134 | 0.29964 | 0.27466 | 0.24987 | 0.23033 | 0.21365 | 0.20113 | 0.19251
Time 17.77 h | 0.0035496 | 0.0040835 | 0.0039521 | 0.0028215 | 0.0042737 | 0.0027541 | 0.0027289 | 0.002973 | 0.003524 | 0.0041116 | 0.0049544 | 0.005915 | 0.0071383 | 0.008657 | 0.010754 | 0.013457 | 0.017098 | 0.021944 | 0.02855 | 0.037527 | 0.049949 | 0.066706 | 0.089582 | 0.12071 | 0.161 | 0.21078 | 0.26707 | 0.32364 | 0.37298 | 0.40854 | 0.42783 | 0.43205 | 0.42611 | 0.41439 | 0.40138 | 0.39048 | 0.38079 | 0.37269 | 0.36524 | 0.35334 | 0.34137 | 0.32206 | 0.30077 | 0.27948 | 0.26059 | 0.24127 | 0.22398 | 0.21022 | 0.20091 | 0.19452 | 0.19064
Time 18.77 h | 0.0030254 | 0.0036888 | 0.0035696 | 0.0025099 | 0.0038466 | 0.0034824 | 0.0022998 | 0.0023773 | 0.0027732 | 0.0031891 | 0.0037944 | 0.0044769 | 0.0053587 | 0.0064883 | 0.008037 | 0.010114 | 0.012984 | 0.016898 | 0.022397 | 0.02989 | 0.040443 | 0.054773 | 0.074419 | 0.10098 | 0.13509 | 0.17669 | 0.2229 | 0.26814 | 0.30632 | 0.33262 | 0.34577 | 0.34724 | 0.34123 | 0.33157 | 0.32171 | 0.31327 | 0.30668 | 0.29974 | 0.29194 | 0.28555 | 0.27137 | 0.25997 | 0.2468 | 0.22718 | 0.21645 | 0.20438 | 0.19349 | 0.18677 | 0.18222 | 0.18195 | 0.18215
Time 19.77 h | 0.0026893 | 0.0045518 | 0.0031729 | 0.0030749 | 0.0033519 | 0.0029816 | 0.0018113 | 0.0017792 | 0.0019425 | 0.0021781 | 0.0025253 | 0.0029148 | 0.0034291 | 0.0040933 | 0.0050704 | 0.0064403 | 0.0084281 | 0.011244 | 0.015313 | 0.02104 | 0.029181 | 0.040274 | 0.055482 | 0.076008 | 0.10225 | 0.13406 | 0.16917 | 0.20331 | 0.23185 | 0.25115 | 0.26054 | 0.2614 | 0.25666 | 0.24992 | 0.24271 | 0.23676 | 0.23245 | 0.22798 | 0.2221 | 0.2177 | 0.21212 | 0.20197 | 0.19302 | 0.18321 | 0.17641 | 0.17082 | 0.16653 | 0.16392 | 0.16497 | 0.16726 | 0.17299
Time 20.77 h | 0.0022289 | 0.0041033 | 0.0036276 | 0.0025911 | 0.0028514 | 0.002468 | 0.00131 | 0.0012066 | 0.0011838 | 0.0012754 | 0.0013852 | 0.0015187 | 0.0017038 | 0.0019642 | 0.0024287 | 0.003144 | 0.0043048 | 0.0060936 | 0.008816 | 0.012789 | 0.018494 | 0.026349 | 0.037079 | 0.051455 | 0.069659 | 0.091561 | 0.11559 | 0.13884 | 0.15818 | 0.17123 | 0.17755 | 0.17809 | 0.17511 | 0.17055 | 0.16586 | 0.1622 | 0.15945 | 0.15653 | 0.15399 | 0.15123 | 0.14769 | 0.14391 | 0.13919 | 0.13652 | 0.13299 | 0.13063 | 0.1298 | 0.13211 | 0.13557 | 0.14049 | 0.14884
Time 21.77 h | 0.0019639 | 0.0030858 | 0.0040378 | 0.0022931 | 0.0025016 | 0.0021092 | 0.0017667 | 0.00081346 | 0.00066381 | 0.00062677 | 0.00055786 | 0.00049798 | 0.00043389 | 0.00040836 | 0.0004829 | 0.0007404 | 0.0013088 | 0.0023585 | 0.0041293 | 0.0067812 | 0.010698 | 0.016063 | 0.023301 | 0.032803 | 0.044561 | 0.058364 | 0.07312 | 0.087078 | 0.098371 | 0.10572 | 0.10895 | 0.10888 | 0.10688 | 0.10404 | 0.10136 | 0.099554 | 0.098232 | 0.097284 | 0.095937 | 0.095231 | 0.093568 | 0.092485 | 0.091345 | 0.092823 | 0.090824 | 0.091408 | 0.093409 | 0.097152 | 0.10185 | 0.10833 | 0.11552
Time 23.77 h | 0.0014883 | 0.0021929 | 0.0021109 | 0.0019847 | 0.0021538 | 0.0017523 | 0.0013903 | 0.00046392 | 0.00018159 | 7.2884e-06 | -0.00023472 | -0.00047996 | -0.00076586 | -0.0010762 | -0.0013679 | -0.0015949 | -0.0016571 | -0.0014385 | -0.00084739 | 0.00024785 | 0.0019441 | 0.0042525 | 0.0072546 | 0.010935 | 0.015153 | 0.019671 | 0.024092 | 0.027859 | 0.030518 | 0.031886 | 0.032062 | 0.031418 | 0.030371 | 0.029274 | 0.028353 | 0.027687 | 0.027241 | 0.026777 | 0.026583 | 0.026424 | 0.026815 | 0.026934 | 0.027208 | 0.027093 | 0.028499 | 0.029847 | 0.032375 | 0.035587 | 0.039182 | 0.042859 | 0.048689
Time 28.77 h | 0.0017085 | 0.0031247 | 0.0023054 | 0.0022179 | 0.0024728 | 0.0020676 | 0.0017334 | 0.00080214 | 0.00068034 | 0.00061049 | 0.00049485 | 0.00037753 | 0.00022966 | 6.4412e-05 | -0.00012393 | -0.00032652 | -0.00051989 | -0.00068838 | -0.0008251 | -0.0009233 | -0.0009599 | -0.0009245 | -0.00086613 | -0.00075777 | -0.00061839 | -0.00048758 | -0.00035858 | -0.00025679 | -0.000189 | -0.00014685 | -0.00010315 | -5.504e-05 | -5.28e-05 | -5.1501e-06 | 6.6811e-05 | 0.00017406 | 0.00023978 | 0.00062246 | 0.00062368 | 0.00065183 | 0.00024422 | 0.00033093 | 0.0019106 | 0.00044643 | 0.00077331 | 0.00090104 | 0.00096653 | 0.0012018 | 0.0015582 | 0.0030805 | 0.002018
10
15
Time / h
6
4
3
2
1
Log Frequency
5
0.0
0.2
0.4
0.6
dielectric loss
-C" /pF
Log Frequency
2
3
4
5
6
1
-42.6oC
-37.1 oC
-32.3oC
-26.7 oC
-30.5 oC
-30.1oC
-29.5 oC
-30.1oC
Primary Drying
20
25
30
35
-25.1 oC
-19.1 oC
Time 12 h
Time 38 h
Primary Drying
Notes: Comment in version 14 – if we are using the proximity of a merged peak to explain why the re-cooling temp calibration gives a poor temp prediction then surely the same would apply here – as the low frequency shoulder is quite significant and yet the agreement between TVIS and the T(Fpeak) determined from the annealing stage is excellent.
Full shelf 5%w/v Lactose in ultrapure water 23Jan17-CH3 (Fill factor 0.7 = 3.04 g)-3mm above base electrode

## Slide 29: C" PEAK Criteria & Assumptions
C" PEAK can be considered in three regions
A linear region (% of C" PEAK est. on water in isolated vial)
A non-linear region
An ice cone region – where C" PEAK ~ 0 (use C’ @ high freq.)
C" PEAK is proportional to the height of the ice cylinder bounded by the electrode region, through the value of CG
Drying rates are based on the assumption of a planar sublimation front
Below the electrode C" PEAK loses sensitivity to ice layer height (non-linear)
C" PEAK cannot be used for end point determination: use C’ @ high freq. instead.
Linear region
Non-linear region
C”PEAK
Height of ice layer, h
Height of ice layer below the electrode, hbe
Height of ice layer at the electrode region, he
80%
Notes: C" PEAK can be considered three regions
A linear region (where C”peak is proportional to the ice within the electrode region),
A non-linear region, where C”peak has a decreasing sensitivity to the ice below the electrode
A region where C”peak might still be sensitive to the ice cone.

## Slide 30: Drying Rate (dm/dt) Determination
Chart: Time/h
Chart: 
| -47.7 | -47 | -46.1 | -45.1 | -44.2 | -43.3 | -42.1 | -41.1 | -40.1 | -38.800000000000011 | -37.6 | -36.5 | -35.4 | -34.1 | -32.800000000000011 | -31.5 | -30.3 | -29.1 | -27.8 | -26.4 | -25.3 | -24 | -22.8 | -21.5 | -20.3
Re-heating | 0.42283 | 0.42509 | 0.4276 | 0.43028 | 0.43297 | 0.43584 | 0.43889 | 0.44212 | 0.44561 | 0.44941 | 0.45317 | 0.45703 | 0.46055 | 0.46386 | 0.46731 | 0.47143 | 0.47626 | 0.48245 | 0.49002 | 0.49911 | 0.50973 | 0.52232 | 0.53734 | 0.55474 | 0.57399
Chart: Temp. /oC
Chart: Ice mass / g
Chart: Time/h
Temperature
correction
Convert to
ice content
Convert to drying rate
Notes: Full shelf 5%w/v Lactose in ultrapure water 23Jan17-CH3 (Fill factor 0.7 = 3.04 g)-3mm above base electrode

## Slide 31
TVIS Application Primary Drying: Rp determination
5%w/v Lactose in deionised water

## Slide 32: Lactose Dried Product Resistance (RP)
Chart: Temp. /oC
Chart: Pressure / torr
9.1 mm
3 mm
Chart: Time/h
Chart: Dry Layer Thickness (cm)
1.75cm2·Torr·h·g-1
Notes: Version 14
If we are now calculating Rp then we only need to show the product temperature (Tfpeak) and the condenser temperature
Then below that the values for partial pressure – to replace the log Fpeak plot
Then the drying rate (ice vs time) and not c”peak vs time on the bottom left
Replace the top right with my new Rp values
Full shelf 5%w/v Lactose in ultrapure water 23Jan17-CH3 (Fill factor 0.7 = 3.04 g)-3mm above base electrode
From the literature (Pikal (1985) PDA Journal of Pharmaceutical Science and Technology, 39(3), pp.115-139.): Rs ~0.33 Torr·h·g-1.
The temperatures we are using (calculated from log Fpeak) have not been smoothed and the values for C"peak (although taken from the moving average) both have scatter to them - hence why the plot of Rp vs dry layer is quite noisy
Rp starts at 2 (which I presume is the stopper resistance) then it increases as the dry layer develops in thickness up to 3 mm - which is the point at which micro-collapse has happened. At that point, the overall resistance decreases which means that microcollapse has also occurred in the layer above the ice - which already has no ice. If there was no microcollapse in that layer then Rp would stay constant after 3 mm of the dry layer develops. Instead, Rp drops to a new value at ~ 4mm dry layer (when microcollapse occurs) and doesn't change after that. I presume that is because a more dramatic microcollapse occurs in the layers that develop below 4 mm - which are seen on our SEM as large pores and so the layers below the top 3 mm do not have a measureable dry layer resistance.
Milton et al. study product resistance by comparing the direct vial method and MTM (data shown in literature are from direct vial technique according to Pikal's method). They concluded that the product resistance value from MTM agree with direct vial technique and has the same pattern. For lactose, they reported that product resistance increases continuously with thickness up to the onset of collapse (below about -30 Celcius).

## Slide 33: Lactose Dried Product Resistance (RP)
Middle layer
Micro-collapse
Top layer
Fine pores
Bottom layer
Full collapse
500X
3 mm
500X
500X
Chart: Dry Layer Thickness (cm)
3 mm
Notes: Full shelf 5%w/v Lactose in ultrapure water 23Jan17-CH3 (Fill factor 0.7 = 3.04 g)-3mm above base electrode
Img: 5%lactose_16Nov16-Ng03-10a

## Slide 34
TVIS Application Primary Drying: Kv determination
Notes: Version 14 – this is a new slide

## Slide 35: Heat Transfer Coefficient (KV) Determination
Chart: Time / h
Chart: Temperature calibration
| 2.6862339999999998 | 2.6881900000000041 | 2.6941359999999999 | 2.704088 | 2.7179570000000002 | 2.735554 | 2.7549359999999998 | 2.7764509999999967 | 2.8000249999999998 | 2.8233779999999999 | 2.8486829999999967 | 2.8758249999999967 | 2.9053 | 2.9347489999999961 | 2.9657499999999977 | 2.9988519999999967 | 3.0339119999999999 | 3.0669900000000001 | 3.1020270000000001 | 3.1388659999999966 | 3.1758619999999977 | 3.2130019999999999 | 3.2517659999999977 | 3.2905300000000035 | 3.329294 | 3.3680579999999987 | 3.4068219999999987 | 3.445586 | 3.4843500000000001 | 3.525433 | 3.5639989999999999
Re-Heating | -41.6 | -41.6333 | -41.4 | -41.0333 | -40.3667 | -39.7667 | -38.8333 | -38.1667 | -37.2333 | -36.4667 | -35.4 | -34.5667 | -33.6 | -32.8 | -31.6667 | -30.7667 | -29.8333 | -28.8333 | -27.9 | -26.7667 | -25.8333 | -24.7667 | -23.8 | -22.7333 | -21.8667 | -20.9 | -19.8667 | -18.9333 | -17.9333 | -16.9667 | -15.9
Chart: Primary drying
The product temperature (), which derived by TVIS as T(FPEAK) or by thermocouple (TC), is one of parameters needed for KV determination
Re-Heating
Calculate from
Log FPEAK (TVIS)
T (FPEAK) −34.2 oC
TC
TSHELF −20 oC
Sublimation rate (dm/dt) is estimated by TVIS
Test sample is 2 mL water filled in 10 mL vial
Drying rate (dm/dt)
~ 0.3 g/h
and dm/dt from TVIS
KV determination
Applying vacuum
TSHELF recovers
Notes: Full Shelf Water Study_010716_Ch1 (Filling volume = 2 mL) electrode touch base
It seems that there is an initial drying rate that projects to an end point of 18 h. So let’s say we have 2 g that would be removed in 7 h is that drying rate continues to the end of the process. Dm/dt = 0.286 g / h

## Slide 36: Heat Transfer Coefficient (KV) Determination
First convert dm/dt to dq/dt using
the latent heat of sublimation
(L = 2844 J g-1 )
dm/dt = 0.3 g/h
L = 2844 J g-1
dq/dt = 853 J/h
Brülls, M., and Ramusson, A. (2002) Heat Transfer in Lyophilization. Int J Pharm 10;246(1-2):1-16.
Kv values for 10 mL tubing vials (2 mL fill volume)
TSHELF = − 20 oC
TPRODUCT = − 34.2 oC
Av = 0.0045 m2
(Schott Type1 glass 10 ml tubing vial )
Kv = 37 W m-2 K-1 [@ 40 Pa, 400 μBar]
Chart: Pressure (Pa)
| 4.6551724137931032 | 15.517241379310345 | 28.706896551724103 | 48.879310344827665 | 89.224137931034448
O.H.T.C (water exp.) | 9.1139 | 18.4304 | 28.3544 | 46.3797 | 75.1392
O.H.T.C (ice exp.) | 8.1013 | 18.4304 | 51.443
Heat radiation (water exp.) | 62.0253 | 34.0253 | 21.9747 | 14.1772 | 8.1519
Heat radiation (ice exp.) | 42 | 18.6076 | 7.0886
Kv from TVIS | 37
Notes: Full Shelf Water Study_010716_Ch1 (Filling volume = 2 mL) electrode touch base
We can conclude that at this pressure (40 Pa) the dominant heater transfer mechanism is vis the base rather than radiation

## Slide 37: Acknowledgements, Recent Projects & Collaborators
De Montfort University
  Evgeny Polygalov. Senior Research Fellow
  Irina Ermolina. Senior Lecturer
  Yowwares Jeerarunangrattana. PhD student
  Bhaskar Pandya. PhD student
GEA Process Engineering
  Trevor Page & Julian Taylor
  Daniela Buchmeyer & Thomas Beutler
BlueFrog : Chris Samwell Ben Irvin
NIBSC : Paul Matejtschuk
Sanofi : Tim McCoy

## Slide 38: Thank you

## Slide 39: Annex : TVIS Theory

## Slide 40: Equivalent circuit model
Ionic diffusion layer (dI )
“Polarization of glass”
“Polarization of glass-solution interface”
“Polarization of solution/ice”
Glass wall
Parallel-plate capacitor equation
CI
dl
Electrode
+
+
Cg
dg
Measurement vial
Electrode
−
Cg
−
CS
CI
+
+
+
+
+
+
+
+
+
−
−
+
−
+
−
+
−
+
dS
dl
dg
Glass wall
−
−
−
−
−
−
−
−
Cg
CS
CI
Cg
CI
RS
+
−
Glass wall
−
−
−
−
−
−
−
+
+
+
+
+
+
+
−
−
−
−
−
−
−
+
+
+
+
+
+
+
δ+
δ−
δ+
δ−
δ+
δ−
δ+
δ−
δ+
δ−
δ+
δ−
+
Notes: Version 11 : add typical values for each capacitance for a water filled vial. Add the resistance values for liquid water and a ice
Can you make the symbols for + and – larger – make sure you use the symbol for the negative sign rather than the hyphen
Add a resistor in parallel with Cs (above the line of capacitors)
Add the equation for the resistor in terms of resistivity and the cell constant (in this case, it will be ds/A)
Add some –ve and +ve charges migrating in opposite directions through the bulk solution (shows wiggly arrows to indicate diffusion/brownian motion) to show the phenomenon of conductivity

## Slide 41: Impedance Spectroscopy
Impedance is a frequency dependent parameter largely because the impedance of a capacitance is dependent on the frequency of the applied field, whereas an ideal resistor has zero frequency dependence
C
R
Θ = 0°
│Z│ = R
Θ = 90°
│Z│ = 1/ωC
│Z│
│Z│
Notes: Use fit data from Full Shelf Water Study_010716_Ch1_45.z (frozen state) C1 = Cs (3.096e-14), R1 = Rs(2.018e8), C2 = CG(6.184e-13)
This slide use CG, Rs
Start with the value for the capacitor as CG and then use that for the series
Edited in version 4
The impedance of most real objects (i.e. those which have a combination of characteristics, for example, some resistance and some electrical capacitance) is a frequency dependent parameter largely because the impedance of a capacitance is dependent on the frequency of the applied field, whereas an ideal resistor has zero frequency dependence
The impedance spectrum of a capacitor )in terms of the impedance magnitude and phase angle) displays a characteristic negative slow of -1 in the plot of log impedance magnitude vs log frequency. This is because the impedance is an inverse function of the capacitance and the applied frequency

## Slide 42: Impedance Spectroscopy
In the case of a composite object that has both capacitance and resistance then the impedance spectrum that results will be dominated by one or the other element.
In a series circuit
@ low frequency the capacitor dominates the spectrum because the impedance of the capacitance is so high that the capacitor effectively controls the current that flows through the circuit
@ high frequency the resistor dominates the spectrum because the impedance of the capacitor has fallen below that of the resistor such that the resistor effectively controls the current that flows through the circuit
C
R
Θ = 45°
│Z│ = R
│Z│ = 1/ωC
│Z│
Notes: Use fit data from Full Shelf Water Study_010716_Ch1_45.z (frozen state) C1 = Cs (3.096e-14), R1 = Rs(2.018e8), C2 = CG(6.184e-13)
This slide use CG, Rs
Edited in version 4
The impedance of most real objects (i.e. those which have a combination of characteristics, for example, some resistance and some electrical capacitance) is a frequency dependent parameter largely because the impedance of a capacitance is dependent on the frequency of the applied field, whereas an ideal resistor has zero frequency dependence
The impedance spectrum of a capacitor (in terms of the impedance magnitude and phase angle) displays a characteristic negative slow of -1 in the plot of log impedance magnitude vs log frequency. This is because the impedance is an inverse function of the capacitance and the applied frequency

## Slide 43: Impedance Spectroscopy
In the case of a composite object that has both capacitance and resistance then the impedance spectrum that results will be dominated by one or the other element.
In a parallel circuit
@ low frequency the resistor dominates the spectrum because the impedance of the capacitance is so high that all the current flows through the resistor.
@ high frequency the capacitor dominates the spectrum because the impedance of the capacitance is now lower than the resistor such that all the current now flows through the capacitor.
R
C
Θ = 45°
│Z│ = R
│Z│
│Z│ = 1/ωC
Notes: Use fit data from Full Shelf Water Study_010716_Ch1_45.z (frozen state) C1 = Cs (3.096e-14), R1 = Rs(2.018e8), C2 = CG(6.184e-13)
This slide use CG, Rs
Edited version 11
Edited in version 4
The impedance of most real objects (i.e. those which have a combination of characteristics, for example, some resistance and some electrical capacitance) is a frequency dependent parameter largely because the impedance of a capacitance is dependent on the frequency of the applied field, whereas an ideal resistor has zero frequency dependence
The impedance spectrum of a capacitor )in terms of the impedance magnitude and phase angle) displays a characteristic negative slow of -1 in the plot of log impedance magnitude vs log frequency. This is because the impedance is an inverse function of the capacitance and the applied frequency

## Slide 44: Impedance Spectroscopy
More complex composite objects can be considered as combinations of impedances. Again, the impedance spectrum that results will be dominated by one or the other impedance.
In a complex circuit
@ low frequency (<1 kHz) the resistor R1 dominates the impedance of the R1C1 circuit, but because this circuit is in series with a capacitor , C2 (which has a high impedance at low frequency) then C2 effectively controls the current that flows through the entire circuit
@ intermediate frequency (1-30 kHz) the impedance of C2 drops below that of the resistor, such that the resistor begins to dominate the impedance and therefore the phase angle tends to increase from -90 to zero
@ high frequency (>30 kHz) the impedance of the capacitor, C1, which is in parallel with the resistor, decreases below that of the resistor such that the resistor no longer dominates the impedance of the parallel RC circuit so then the circuit behaves like two capacitors in series but with C1 dominating the spectrum
C2
R1
C1
1 kHz
30 kHz
│Z│
│Z│ = R
│Z│ = 1/ωC(1)
│Z│ = 1/ωC(2)
Notes: Use fit data from Full Shelf Water Study_010716_Ch1_45.z (frozen state) C1 = Cs (3.096e-14), R1 = Rs(2.018e8), C2 = CG(6.184e-13)
Version 11 : what ever value you have used for C2 – can you use that for the Capacitor in the series circuit and the parallel circuit then create another capacitor for C1 which gives a response like you have currently when your original C1 and C2 are placed in series. That way we can have the capacitor line 2 the same on RC series and RC parallel
It looks like there is ~10 fold difference in the two capacitors that you’ve chosen. Is that realistic in terms of the capacitances of the interface and the solution?
Edited in version 4
The impedance of most real objects (i.e. those which have a combination of characteristics, for example, some resistance and some electrical capacitance) is a frequency dependent parameter largely because the impedance of a capacitance is dependent on the frequency of the applied field, whereas an ideal resistor has zero frequency dependence
The impedance spectrum of a capacitor )in terms of the impedance magnitude and phase angle) displays a characteristic negative slow of -1 in the plot of log impedance magnitude vs log frequency. This is because the impedance is an inverse function of the capacitance and the applied frequency

## Slide 45: Impedance Spectroscopy
The impedance spectrum of complex element can be presented as the capacitance spectrum
C2
R1
C1
1 kHz
30 kHz
│Z│
│Z│ = R
│Z│ = 1/ωC(1)
│Z│ = 1/ωC(2)
Impedance Spectrum
Capacitance Spectrum
Notes: Use fit data from Full Shelf Water Study_010716_Ch1_45.z (frozen state) C1 = Cs (3.096e-14), R1 = Rs(2.018e8), C2 = CG(6.184e-13)

## Slide 46: Interfacial Polarization Characteristic
At
As the frequency increase, increases to maximum ()then decreases to zero as the frequency
At a frequency of
Chart: Log Frequency

## Slide 47: Interfacial Polarization Characteristic
The impedance of the model can be calculated from the following equation
which re-arranges to
Impedance can be expressed in terms of a complex capacitance
The complex capacitance can also be expressed in form of real part and imaginary part
From the complex capacitance formula, the expressions for real and imaginary capacitance can be calculated to explain the origin of interfacial polarization peak. This achieved by multiplying the nominator and denominator by the complex conjugate of the denominator and by grouping the real () and imaginary () parts
and