## Slide 1: Freeze-drying process optimization using through vial impedance spectroscopy EPSRC EHDA Network Intl. Pharm. Tech. Conference, Leicester Friday 4th November, 2016
Prof. Geoff Smith
Leicester School of Pharmacy

## Slide 2: Outline
Freeze-drying and the design space
Introduction of Through Vial Impedance Spectroscopy (TVIS) Technology
TVIS Applications for Determination of
  Heat transfer coefficient through the vial base KV
  Dry layer resistance which
  impedes the of water vapour sublimation RP
  RP
  Kv

## Slide 3
Biopharma Technology Ltd (BTL),n.d.
Freeze-drying steps: freezing, annealing, primary drying, secondary drying
Lyophillization Steps
Product freezes
Shelf temperature reduced
Annealing of product to optimize ice formation
Apply vacuum
Drying end point
4. Secondary drying
Must happen below collapse temperature, e.g. Tc = -31.4 C

## Slide 4: Primary Drying
The sublimation of ice from the frozen product to achieve a dried layer of solute
The product temperature should be lower than Tc
Schwegman, 2011
Heat and mass transfer in primary drying

## Slide 5: Design Space for Primary Drying

## Slide 6: Design Space
10°
Shelf Temperature
Chamber Pressure (Pc) 50 – 300 mTorr
Product temperature (Tp)
-40 °
-35 °
-30 °
-25 °
-20 °
Sublimation Rate (dm/dt)
Use a range of
chamber pressures at a
fixed shelf temperature (equipment settings)
to establish the relationships between sublimation rate and the product temperature
Notes: The design space is developed as an envelope in a graph of sublimation rate, shelf temperature, and chamber pressure. One boundary of the design space is established by failure of the formulation under aggressive cycle conditions of elevated temperature – causing the product to collapse. Other boundaries are established by equipment performance, inc. Refrigeration performance condenser capacity, heating capacity or limitations in the dynamics of water vapour flow within the system

## Slide 7: Design Space
-40 °
-35 °
-30 °
-25 °
30°
20°
0°
-10°
10°
-20 °
Shelf Temperature
-40 °
-30 °
-25 °
-20 °
-35 °
-40 °
-30 °
-25 °
-35 °
Product temperature (Tp)
Chamber Pressure (Pc) 50 – 300 mTorr
Sublimation Rate (dm/dt)
Repeat at different shelf temperatures
Notes: The design space is developed as an envelope in a graph of sublimation rate, shelf temperature, and chamber pressure. One boundary of the design space is established by failure of the formulation under aggressive cycle conditions of elevated temperature – causing the product to collapse. Other boundaries are established by equipment performance, inc. Refrigeration performance condenser capacity, heating capacity or limitations in the dynamics of water vapour flow within the system

## Slide 8: Design Space
-40 °
-35 °
-30 °
-25 °
30°
20°
0°
-10°
10°
Equipment constraints
Product collapse
-20 °
Shelf Temperature
-40 °
-30 °
-25 °
-20 °
-35 °
-40 °
-30 °
-25 °
-35 °
Product temperature (Tp)
Chamber Pressure (Pc) 50 – 300 mTorr
Sublimation Rate (dm/dt)
And then map on the
equipment constraints (chocked flow) (Pi = Pc )
and the product constraint (collapse) Ti = Tc
To complete the design space (as shown by the yellow triangle)
Then operate at the apex of the triangle to drive process efficiencies
Notes: The design space is developed as an envelope in a graph of sublimation rate, shelf temperature, and chamber pressure. One boundary of the design space is established by failure of the formulation under aggressive cycle conditions of elevated temperature – causing the product to collapse. Other boundaries are established by equipment performance, inc. Refrigeration performance condenser capacity, heating capacity or limitations in the dynamics of water vapour flow within the system

## Slide 9: Through Vial Impedance Spectroscopy (TVIS)

## Slide 10: Recent Funding
LyoDEA
£217 160 Collaborative R&D funding
(Nov ’08-Oct ’12) Innovate UK
BioStaRT
£367 567 Collaborative R&D funding
(Aug ’14- Jul ’17) Innovate UK
AtlasBio
£803 846 Collaborative R&D funding
(Oct ’16-Sept 18) Innovate UK
IP: GB2480299
Electrical monitoring of a lyophilization process
Priority Date: 12th May 2010
Assignee: GEA Pharma Systems
AtlasBio

## Slide 11: Through Vial Impedance Spectroscopy (TVIS)
Impedance measurements across a vial rather than within the vial
Hence “Through Vial Impedance Spectroscopy”
Features
Single vial “Non-product invasive”
Both freezing and drying characterised in a single technique
Stoppering still possible and on-perturbing to the packing of vials
Mesoscale of unfrozen fraction accessible by assessing the temperature dependence of impedance
On going TVIS development for multiple scales with top down electrodes:
  Product Development (Microplate, Micro-vial, or Single vial)
  Mini-Pilot (Small population clusters of vials)
  Scale up to Batch (Large population cluster of vials)

## Slide 12: Through Vial Impedance Spectroscopy (TVIS)
Freeze Dryer and TVIS System
Virtis Advantage Pro
Freeze Dryer
TVIS System
Impedance test vial
Pass-through
Junction box

## Slide 13
Imaginary Part of Capacitance
Real Part of Capacitance
High frequency
Liquid state
Liquid state
Frozen solid
Frozen solid
low frequency
Annealing = Re-heating and Re-cooling
Re-heating
Re-heating
Intermediate frequency
Re-cooling
Re-cooling
low frequency
Primary drying
Primary drying
low frequency
TVIS Response Surface

## Slide 14: TVIS Response Surface
Freezing Process
Primary Drying Process
Freezing
Freezing
11 h
12 h
13 h
15 h
14 h
16 h
17 h
18 h
19 h
11 h
12 h
13 h
15 h
14 h
17 h
16 h
18 h
19 h
24.1 oC
12 oC
-5.7 oC
-24.3 oC
-24.3 oC
-5.7 oC
12 oC
24.1 oC

## Slide 15
The magnitudes of CG and CS are proportional to the area of interface between the frozen mass and the glass adjacent to the electrode ().
𝐶𝐺 = A
𝐶S = A
𝐴 = A1
dS
Constant (KG)
Constant (KS)
Primary Drying
𝐴 = A2
Ice sublime
1
2
dS
In case of a flat sublimation interface, the interfacial area between the frozen layer and the juxtaposed glass wall () will decrease in proportion to the remaining ice volume
dG
dG

## Slide 16: TVIS Technology: Principle
The capacitance spectrum is dependent on both the electrical resistance and electrical capacitance of the vial contents.
  Data viewing software (LyoView ™) identifies the peak frequency (FPEAK) and peak amplitude (CʺPEAK) in the imaginary part of the capacitance spectrum
In general terms:
  FPEAK can be used to monitor phase behaviour (ice formation, glass transitions) and product temperature
  CʺPEAK can be used to monitor the amount of ice remaining during primary drying, from which the drying rate and end point may be determined.

## Slide 17: TVIS Application Heat Transfer Coefficient (KV) Determination (Test sample: 2 mL water in 10 mL vial)

## Slide 18: Primary Drying Modelling : Heat and Mass Transfer
Dry layer resistance Rp
Thickness of Ice Layer (dICE)
Chamber pressure (PCHAMBER)
TPRODUCT & PICE
Internal cross sectional area of one vial (AV)
ΔT
Shelf TSHELF
KV determination
dm/dt - sublimation rate (g/hour/vial)
PICE - Vapour pressure at ice sublimation interface (Torr)
PCHAMBER - Chamber pressure (Torr)
- Area normalized product and stopper resistance
dq/dt - heat flow to the product
L - Latent heat of sublimation (J g-1)
AV - cross-sectional area of the vial (cm2)
TSHELF - Shelf temperature (K)
- Product temperature at the bottom of vial (K)
dICE - Thickness of the ice layer calculated (cm)
∆T - Temperature different between ice sublimation
front and the bottom of the vial (K)
TANG, X.C., NAIL, S.L. and PIKAL, M.J. (2005) Freeze-drying process design by manometric temperature measurement: design of a smart freeze-dryer. Pharmaceutical Research, 22 (4), pp. 685-700.
Stopper resistance Rs
Notes: Pice Vapour pressure at ice sublimation interface (determined from fit Torr)
Po Chamber pressure (measured, Torr)
V is volume of freeze dryer chamber (Known, L)
N Total No. of filled vials (Know, No.)
Ap Inner cros-sectional area of the vial (Known, cm2)
Ts Shelf temperature (set, K)
Rp, Rs area normalized product and stopper Resistance
Lice Thickness of the ice layer calculated (cm)
∆T temperature diff b/w ice sublimation front and the bottom of the vial
X is constant (Torr/sec)

## Slide 19: Heat Transfer Coefficient (KV) Determination
Chart: Primary drying
The product temperature (), which derived by TVIS as T(FPEAK) or by thermocouple (TC), is one of parameters needed for KV determination
Chart: Temperature Calibration
Re-Cooling
Calculate from Log FPEAK (TVIS)
T (FPEAK)
TC
TSHELF
Sublimation rate (dm/dt) is estimated by TVIS
Test sample is 2 mL water filled in 10 mL vial
Drying rate (dm/dt)
~ 0.3 g/h
and dm/dt from TVIS
KV determination
Chart: Time / h
Applying vacuum
− 34.2 oC
TSHELF recovers
− 20 oC
Notes: It seems that there is an initial drying rate that projects to an end point of 18 h. So let’s say we have 2 g that would be removed in 7 h is that drying rate continues to the end of the process. Dm/dt = 0.286 g / h

## Slide 20: Heat Transfer Coefficient (KV) Determination
First convert dm/dt to dq/dt using
the latent heat of sublimation
(L = 2844 J g-1 )
dm/dt = 0.3 g/h
L = 2844 J g-1
dq/dt = 853 J/h
Brülls, M., and Ramusson, A. (2002) Heat Transfer in Lyophilization. Int J Pharm 10;246(1-2):1-16.
Kv values for 10 mL tubing vials (2 mL fill volume)
TSHELF = − 20 oC
TPRODUCT = − 34.2 oC
Av = 0.0045 m2
(Schott Type1 glass 10 ml tubing vial )
Kv = 37 W m-2 K-1 [@ 40 Pa, 400 μBar]
Chart: Pressure (Pa)
| 4.6551724137931032 | 15.517241379310345 | 28.706896551724135 | 48.879310344827594 | 89.224137931034477
O.H.T.C (water exp.) | 9.1139 | 18.4304 | 28.3544 | 46.3797 | 75.1392
O.H.T.C (ice exp.) | 8.1013 | 18.4304 | 51.443
Heat radiation (water exp.) | 62.0253 | 34.0253 | 21.9747 | 14.1772 | 8.1519
Heat radiation (ice exp.) | 42 | 18.6076 | 7.0886
Kv from TVIS | 37
Notes: We can conclude that at this pressure (40 Pa) the dominant heater transfer mechanism is vis the base rather than radiation

## Slide 21: TVIS Application Product Resistance (Rp ) Determination (Test sample: 2 mL protein solution fill in 10 mL vial)

## Slide 22: Dried Product Resistance (RP) Determination of Protein Solution
A 50% v/v protein solution 2 mL has protein 0.042 g
Protein 1 g has ~0.3 g (unfrozen water) unavailable water (Pang 2014)
Therefore unfrozen water of protein is 0.3 x 0.042 = 0.0126 g
The weight of ice is approximately 2 – 0.042 - 0.0126 = 1.932 g
PBS
4.2% w/v Protein Solution
16 mL
16 mL
Preparation: 50% v/v Protein Solution
50% v/v Protein Solution
2 mL
Chart: Time / h
Freeze drying Protocol
Apply Vacuum
Pang, X.-F. (2014) Water: Molecular structure and properties. Singapore, Singapore: World Scientific Publishing Co Pte.
Notes: Weight of ice calculated based on a 4.2% Albumin in protein solution and albumin 1 g has unavailable water to colligative effect 0.3 g (unfrozen water)
2 mL of 50% v/v protein solution has albumin 0.042 g and unfrozen water 0.3 x 0.042 = 0.0126 g
So, weight of ice = 2-0.0126 = 1.98 g
Drying rate : 1.98 g in 30 – 10.6 h

## Slide 23: Product characterization: Opportunity for Protein Solution
Low frequency process may hold information on the state/properties of the protein layer, inc. progression and end point of secondary drying
Low frequency process
(end of primary drying)
High frequency process
(start of primary drying)
Ice layer
protein layer

## Slide 24: Dried Product Resistance (RP) Determination of Protein Solution
Chart: Time / h
Chart: Time / h
Chart: Time / h
The peak at high frequency (~12 kHz) is due to the ice layer
  Ice layer peak decreases during the early stages of primary drying (< 25 h) because the peak amplitude is proportional to the volume of the ice layer
The peak at low frequency is due to the dry layer
    Dry layer peak increases during the later stages of primary drying (> 27h) because the peak amplitude is proportional to the volume of the dry layer
11 h
13 h
15 h
17 h
19 h
21 h
23 h
25 h
27 h
29 h
31 h
33 h
35 h
Primary drying process
(-48.9 to – 42.1 oC)
Frozen state (-41.8oC)
12 kHz
TSHELF
T(FPEAK )
TAVERAGE
Calibration data for log FPEAK during annealing is used to predict the product temperature (T(FPEAK)
48.1 oC (3h)

## Slide 25: Dried Product Resistance (RP) Determination of Protein Solution
ICE
Shelf
Dried Product (Rp)
ICE
Dried Product
Water vapor
Heat
Ice-product interface
The different types of resistance to mass transport (water vapour flow)
  dm/dt = the rate of mass transfer for the water vapour or sublimation rate (g/hour/vial)
  PICE = the equilibrium vapour pressure of ice at the sublimation interface temperature (Torr)
PCHAMBER = the chamber pressure (Torr)
  = the area normalized resistance of the dried product (cm2·Torr·h·g-1)
  Ap = the cross-sectional area of the
product (cm2)
Dried product resistance
Semi-stoppered resistance
Chamber to condenser resistance
Condenser
Sublimation rate
TANG, X.C., NAIL, S.L. and PIKAL, M.J. (2005) Freeze-drying process design by manometric temperature measurement: design of a smart freeze-dryer. Pharmaceutical Research, 22 (4), pp. 685-700.
Product resistance (RP) is ~ 80% of the total mass transfer resistance (Pikal, Roy and Shah (1984).
Notes: Define Rs and NRC
Does NRC relate to the duct between the chamber and the condensor – if so then we can rule that one out.
And why is it not just RsubscriptC
PC should have subscript c
Can you animate the equation so that the top one disappears when you show that R(tinhat)p is in fact a combination of all three resistances
Can you check the literature for definition of Pc : is it chamber or condenser? And do they mean total pressure or partial pressure of ice. If the latter then they are doing the same as us, ie using the condenser to calculate the partial pressue of ice in the chamber (or more specifically at the surface of the condensor) - given that it is inside the chamber.
In the process, can you also check what we mean by chamber pressure – what gauge is used on our system? I think it measures total pressure not just the partial ice vapour pressure – which is the one that’s important in driving the process.

## Slide 26: Dried Product Resistance (RP) Determination of Protein Solution
can estimate by TVIS
Ice vapour pressure calculated from TPRODUCT the product temperature (as derived by TVIS or Thermocouple)
Chamber pressure calculated from the temperature of condenser (TCONDENSER) (in case of condenser is in the freeze dry chamber)
Cross-sectional area of product is calculated by using internal diameter of test vial
For example, Schott Type1 glass 10 ml tubing vial has internal diameter 2.2 cm therefore, = 3.8 cm2

## Slide 27: Dried Product Resistance (RP) Determination of Protein Solution
Chart: Time / h
The drying rate (dm/dt) estimated by TVIS is one of parameters used for determination of dried product resistance (RP)
First convert C"PEAK from TVIS to Ice mass
C"PEAK is proportional to the amount of ice
Then, calculate drying rate (dm/dt) using time derivative of ice mass
y = 0.001x2 - 0.107x + 2.9457
↓
dm/dt = 0.002X - 0.107
RP determination
0.657 pF ≡ 1.98 g water
Chart: Time / h
Chart: Time / h
1.93 g water

## Slide 28: Dried Product Resistance (RP) Determination of Protein Solution
Chart: Time / h
Ap = 3.8 cm2
(Schott Type1 glass 10 ml tubing vial)
TCONDENSER
TSHELF
Chart: Time / h
TSHELF
TAVERAGE
TCONDENSER
dm/dt
dm/dt
Notes: Add the polynomial for Taverage

## Slide 29: Dried Product Resistance (RP) Determination of Protein Solution
Chart: Dry Layer Thickness (cm)
2% BSA annealing 1 h (Ref. 3) | 56.9207 | 46.1058 | 36.5919 | 32.7294 | 31.1202 | 29.7457 | 28.7697 | 27.805 | 27.9429 | 26.888 | 26.0084 | 25.6143 | 25.1465 | 23.8629 | 23.636 | 22.7127 | 22.2117 | 21.5748 | 21.0214 | 21.201 | 20.5754 | 19.611 | 19.2495 | 19.0984 | 18.5758 | 17.8034 | 17.3933 | 17.01 | 16.5399 | 15.9933 | 15.5515 | 15.2495 | 14.6948 | 14.6206 | 13.9714 | 13.809 | 13.6574 | 13.1687 | 12.8072 | 12.6577 | 12.465 | 12.0595 | 11.8873 | 11.6471 | 11.3233 | 11.0546 | 11.1232 | 10.514 | 10.3565 | 10.2964 | 10.0212 | 10.0069
2% BSA non annealing (Ref. 4) | 55.1724 | 40 | 36.0743 | 31.643 | 29.6552 | 28.1859 | 27.5862 | 26.635 | 25.0398 | 24.7876 | 24.2311 | 23.6155
2% IgG non annealing (Ref. 4) | 50.5747 | 37.7495 | 30.7692 | 29.3103 | 27.5862 | 25.7053 | 24.8817 | 23.6453 | 22.6919 | 21.501 | 20.4062 | 19.904 | 19.7044 | 19.1314
2% Lysozyme non annealing (Ref. 4) | 34.5213 | 29.2553 | 23.2542 | 21.9415 | 19.9468 | 19.0603 | 17.5532 | 16.9187 | 15.8996 | 15.2128 | 14.6277 | 14.2526 | 13.6681 | 13.1979 | 12.864 | 12.4875 | 12.1897 | 11.8448 | 11.4626 | 11.1138 | 10.9157
0.5% BSA annealing 6 h (Ref. 3) | 47.028 | 31.6719 | 26.1688 | 22.0444 | 20.0488 | 18.5974 | 16.952 | 16.6278 | 15.6264 | 15.016 | 14.2585 | 13.7557 | 13.1241 | 12.5749 | 12.1262 | 11.3953 | 11.1519 | 10.6543 | 10.3462 | 9.8983 | 9.6621 | 9.129 | 9.0612 | 8.6938 | 8.5393 | 8.215 | 7.9878 | 7.8604 | 7.5441 | 7.4338 | 7.3609 | 7.1738 | 7.0397 | 7.1825
5% Glycine non annealing (Ref. 2) | 39.96 | 15.6477 | 12.3592 | 11.025 | 10.442 | 10 | 9.72 | 9.4639 | 9.3462 | 9.1703 | 9.1154 | 9.439 | 9.9296 | 10.747
0.5% BSA annealing 1 h (Ref. 3) | 28.2168 | 20.7477 | 16.1239 | 13.9499 | 12.0929 | 11.757 | 10.5813 | 9.8223 | 9.2712 | 8.8639 | 8.3979 | 8.3396 | 7.936 | 7.5968 | 7.521 | 7.1029 | 6.8266 | 6.7913 | 6.5563 | 6.4382 | 6.2916 | 6.2297 | 6.1724 | 6.0513 | 5.9058 | 5.8653 | 5.766 | 5.739 | 5.7243 | 5.7844
1. GIESELER, H., KRAMER, T. and PIKAL, M.J. (2007); 2. JOHNSON, R.E. et al. (2010); 3. LEWIS, L.M. et al (2010)
Rp based on TC temperature
Rp assuming constant temperature (48 oC)

## Slide 30: Summary & Future Work
Kv value determination possible.
Increase the drying time over which T(Feak) maybe determined by increasing the fill volume of ice.
Investigate the dependence of Kv son chamber pressure for design space determination
Rp value determination possible:
Investigate uncertainties in temperature measurement to provide for more reliable estimates of Rp (compare RTDs vs TCs)
Calculate Product Temperature for heat flux (dq/dt) and Kv
Use equivalent circuit modelling to fit spectra and develop a TVIS surrogate temperature model
Characterise a wide range of materials, formulations and process parameters, inc. nucleation temperature, fill volume, freezing rate, annealing

## Slide 31: Summary of Applications of TVIS
Summary of Applications of TVIS
TVIS registers thermal events through changes in the sample resistance associated with
  Discontinuous changes in viscosity (glass transition, collapse)
  Change of state (e.g. ice formation and eutectic formation)
Temperature control might be possible through improved modelling and calibration (Equivalent circuits)
Primary drying (loss of ice) is monitored through changes in the strength of the dielectric loss peak (or step in the real part capacitance)
  Enables drying rate determinations for Kv and Rp calculations
Meso-structural information extracted through the (non-Arrhenius) temperature dependence of the resistance
Mechanisms of annealing may be elucidated from changes in resistance with time (during the heating-hold phase) and from the absence of any changes in TG
Future (with non-contact system)
  Opportunities to track the physical characteristics over a range of scales, from micro-titre plates to collections of vials

## Slide 32: Acknowledgements
GEA Pharma Systems (Trevor Page, Julian Taylor)
BlueFrog Design (Chris Samwell, Ben Irvine, John Hall)
NIBSC (Paul Matejtschuk)
OnkoLytika Ltd (Mark Ecclestone, Annette Williams)
Genzyme Ireland (Tim McCoy)
Innovate UK for part funding this study
DE MONTFORT UNIVERSITY
Evgeny Polygalov. Senior Research Fellow
Yowwares Jeeraruangrattana. Graduate Student
Dr Irina Ermolina. Senior Lecturer

## Slide 33
Any Questions?

## Slide 34: TVIS Application Other Applications

## Slide 35: Through Vial Impedance Spectroscopy (TVIS)
Electrodes are attached to the external surface of a vial
The vial is filled with the sample to fill factor of 0.9
The sample has both resistive & capacitive properties
The glass vial has both resistive & capacitive properties
But glass resistance is so high therefore it
behaves primarily as a capacitor.
In freeze drying, the sample is frozen
and dried by sublimation to obtain a ‘dry’ layer
The dry layer is predominantly capacitive due
to its low moisture content / high resistance
  Exception - Proteins
1 cm
0.9 cm
CS
RS
Glass wall thickness
CS
RS
CG
CS(f)
RS(f)
CS(d)
RS(d)
CS(d)
CS(f)
RS(f)
CS(d)
Cg(d)
Cg(f)

## Slide 36: Through Vial Impedance Spectroscopy (TVIS)
Impedance is a frequency dependent parameter largely because the impedance of a capacitance is dependent on the frequency of the applied field, whereas an ideal resistor has zero frequency dependence
By fitting the impedance spectrum of a composite object then one can extract the sample resistance and capacitance
R
C
R
C

## Slide 37: Imaginary Part Capacitance : Characteristic Response
At
As the frequency increase, increases to maximum ( then decreases to 0 as the frequency
At a frequency of
in radians
in cycles per second

## Slide 38: Product characterization
Chart: Time / h
A
B
D
C
E
Log FPEAK profile of freezing step shows the ice nucleation process:
  Solution is cooled at
  Solution super-cools and starts ice nucleation at This points is referred as “onset of ice nucleation”
  Crystallization or ice nucleation formation occurs during period and continues until the end point of solidification at
A
B
C
D

## Slide 39: Product characterization: Glass transition, Eutectic Crystallization, Collapse
Chart: Temperature / oC
| -27 | -26 | -25 | -25 | -25 | -24 | -23 | -23 | -22 | -22 | -20 | -20 | -19 | -18 | -18 | -17 | -17 | -16 | -15 | -15 | -14 | -13 | -13
 | -4.3814 | -4.3644 | -4.3729 | -4.3684 | -4.3503 | -4.3542 | -4.3506 | -4.3446 | -4.3332 | -4.3246 | -4.3467 | -4.3996 | -4.5425 | -4.8233 | -5.1902 | -5.6349 | -6.0896 | -6.5304 | -7.0663 | -7.6501 | -8.1912 | -8.6156 | -8.8288
 | -4.8233 | -5.1902 | -5.6349 | -6.0896 | -6.5304 | -7.0663 | -7.6501 | -8.1912 | -8.6156 | -8.8288
 | -4.3814 | -4.3644 | -4.3729 | -4.3684 | -4.3503 | -4.3542 | -4.3506 | -4.3446 | -4.3332 | -4.3246 | -4.3467 | -4.3996 | -4.5425
Chart: Time /h
| 1 | 1.0249999999999975 | 1.05 | 1.075 | 1.1000000000000001 | 1.125 | 1.1499999999999975 | 1.175 | 1.2 | 1.2249999999999974 | 1.25 | 1.2749999999999975 | 1.3 | 1.325 | 1.35 | 1.375 | 1.4 | 1.4249999999999972 | 1.45 | 1.4749999999999974 | 1.5 | 1.5249999999999975 | 1.55 | 1.575 | 1.6 | 1.625 | 1.6500000000000001 | 1.675 | 1.7 | 1.7249999999999974 | 1.75 | 1.7749999999999975 | 1.8 | 1.825 | 1.85 | 1.875 | 1.9000000000000001 | 1.925 | 1.9500000000000022 | 1.9750000000000001 | 2 | 2.0249999999999999 | 2.0499999999999998 | 2.0749999999999997 | 2.1 | 2.125 | 2.15 | 2.1749999999999998 | 2.2000000000000002 | 2.2250000000000001 | 2.25 | 2.2749999999999999 | 2.2999999999999998 | 2.3249999999999997 | 2.3499999999999988 | 2.3749999999999987 | 2.4249999999999998 | 2.4499999999999997 | 2.4749999999999988
S5 | 2.684 | 2.868 | 2.988 | 2.972 | 3.292 | 3.208 | 3.56 | 3.56 | 4.128 | 4.24 | 4.832 | 4.884 | 5.488 | -22.252 | 4.704 | 5.564 | 7.46 | 9.148 | 12.396 | 16.412 | 21.74 | 29.08 | 40.8 | 61.06 | 112.412 | 251.68 | 626.4 | 810 | 763 | 654.04 | 528.08 | 392 | 342.8 | 244 | 201.6 | 177.2 | 121.6 | 122.8 | 92 | 56.8 | 62.4 | 25.6 | 58.8 | 34.8 | 8.8 | 59.2 | 28.4 | 33.2 | 47.2 | 8.8 | 49.6 | 22 | -2 | 53.6 | 26 | 31.6 | 35.2 | 12
M5S5 | 2.24 | 2.356 | 2.416 | 2.336 | 2.516 | 2.416 | 2.744 | 2.732 | 3.068 | 3.084 | 3.448 | 3.64 | 3.916 | 4.3 | 4.468 | 1.868 | -12.6 | 6.752 | 9.26 | 12.18 | 15.952 | 20.78 | 27.204 | 36.628 | 52.216 | 80.656 | 137.648 | 271.04 | 468.72 | 600.84 | 638.72 | 573.72 | 496.76 | 400.56 | 316.32 | 258.4 | 199.6 | 169.6 | 136 | 90.8 | 84.4 | 55.2 | 57.6 | 52 | 23.6 | 44.4 | 37.6 | 28.4 | 42.4 | 21.2 | 38 | 32 | 5.2 | 39.6 | 24 | 22 | 37.6 | 20.8
M5% | 0.8248 | 0.8068 | 0.8224 | 0.8192 | 0.8108 | 0.868 | 0.948 | 1.0156 | 1.0796 | 1.142 | 1.2228 | 1.2708 | 1.4564 | 1.5992 | 1.6764 | 1.76 | 1.84 | 1.964 | 2.052 | -9.008 | 2.136 | 3.336 | 4.416 | 5.656 | 7.176 | 9.176 | 11.888 | 16.376 | 24.604 | 41.068 | 77.216 | 175.94 | 395.68 | 570.56 | 631 | 596.84 | 510.2 | 415.36 | 329.68 | 256.64 | 210.8 | 174.8 | 155.2 | 167.6 | 175.2 | 201.6 | 212.8 | 177.6 | 135.6 | 95.6 | 62 | 49.2 | 18 | 15.6 | 14.8 | 16.8 | 22.8 | 14.4
Chart: Time /h
Smith G et al (2013) Eur J Pharm Biopharm 85 (3)1130-40 Arshad MS et al (2014) Eur J Pharm Biopharm 87(3):598-605
Smith G et al (2014) Pharmaceutical Technology 38(4)
Glass transition of 10% maltodextrin DE16-19
Suppression of mannitol eutectic by sucrose
A suddenly change in the capacitance demonstrates the collapse of cake
Notes: DSC measurements (-17 °C)
One of our targets should be to establish the relationship bw Tg and water content for a few polymers and then measure Tg in the vial and work out the water content of the unfrozen mass for different freezing conditions (then see the impact on drying rates – and examine the cake microstructure) the work on drying rates needs to be optimised first to see how fill height impacts the dc/dt curve
In dsc must make sure we measure tg’ – max concentrated solution
Could laser speckle determine the glass transition temperature and thereby enable control of the system close to micro-collapse? Thin collapsed films may be more stable than expanded, porous films.