## Slide 1
1. INTRODUCTION
State-of-the art freeze-drying process control is limited to:
the measurement of pressure (to monitor the duration of primary drying in which ice is removed directly to vapor);
the control of the shelf coolant temperature;
the occasional in-vial temperature measurement (to measure the onset of ice formation) [1].
However, pressure measurements provide only the average drying time of the whole batch and individual vial data is inaccessible.
Moreover, the inclusion of an “in-vial” sensor impacts both the on-set temperature (one is trying to measure) and the structure of ice (which in turn impacts the drying time).
More recent attempts to develop on-invasive methods for measuring individual vial drying times have been proposed. However, these necessitate the application of a mathematical function to the measurement of some in-line physical parameter (e.g. external vial temperature), a process which is non-trivial [2].
2. AIM
The aim of this work is to develop an individual vial monitoring system (LyoSense™), to monitor product characteristics during freeze-drying.
LyoSense™ is a Process Analytical Technology based on the impedance analysis of the lyophilization process.
The electrode array in LyoSense™ is located external to the vial and has low thermal mass
3. OBJECTIVES
To research the application of LyoSense™ technology for the ‘visualization’ of critical events (e.g. product collapse) & cycle phase endpoints (i.e. freezing and primary drying) to establish and control lyophilization regimes.
4. THEORETICAL BASIS OF LYOSENSETM
The LyoSense™ system measures the pseudo-relaxation process associated with the interfacial polarization of the glass wall of the freeze-drying vial, through the resistance of the sample.
The model represented by equivalent circuit shown below satisfactory describes the behaviour of the sample in a freeze drying vial in a course of freeze drying cycle,
Figure 1 Equivalent circuit model for a sample contained within a glass vial.
where C0 is the capacitance formed by the electrodes and glass walls of a vial, C1 is the capacitance of internal volume of a vial and R1 is the electrical resistance of a sample.
This pseudo-relaxation process undergoes characteristic changes in amplitude and peak frequency, such that the progress of freeze-drying and condition of the sample can be controlled (Figure 2).
Figure 2 Illustration of the dependence of the pseudo-relaxation process during the first hours of the pre-freezing stage in the lyophilization of lactose (3% w/v). The shift in the peak frequency, between time points 1h and 1.5 h is indicative of the liquid water to ice phase transition.
5. FREEZE-DRYING METHODOLOGY
Freeze-drying runs were preformed in a 5-shelf HETO FD8 freeze-drier. Vials were arranged in hexagonal arrays (11x 14). A typical cycle comprises: pre-freezing at -30 °C for 2 h; hold for 3h; annealing for 3 h at -10 ° C; primary drying at -20 °C for 15 h; secondary drying for 6h at +20/+40 °C
Chart: log Frequency
| 6 | 5.9030899869919438 | 5.8239109124112867 | 5.6989700043360099 | 5.6020599913279625 | 5.4881133818002974 | 5.3979400086720375 | 5.301029995663991 | 5.204119982655925 | 5.0969100130080545 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521985 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.301029995663991 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.9004927554293745 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463871 | 3.4001233554370853 | 3.3001605369513567 | 3.2001388538573878 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.800036241827986 | 2.7000023971422542 | 2.6000249038929892 | 2.5000030680516976 | 2.4000023459279602 | 2.3000082025538129 | 2.2000018654066062 | 2.0999912335446838 | 2.0074490444977489 | 1.9404816204440281 | 1.8313385741140977 | 1.7064019721397619 | 1.6094876898532873 | 1.506829617234553 | 1.4053634378769695 | 1.308457694189304 | 1.2057996215705744 | 1.1043505192427383 | 1.0002604985473884 | 6 | 5.9030899869919438 | 5.8239109124112867 | 5.6989700043360099 | 5.6020599913279625 | 5.4881133818002974 | 5.3979400086720375 | 5.301029995663991 | 5.204119982655925 | 5.0969100130080545 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521985 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.301029995663991 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.9004927554293745 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463871 | 3.4001233554370853 | 3.3001605369513567 | 3.2001388538573878 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.800036241827986 | 2.7000023971422542 | 2.6000249038929892 | 2.5000030680516976 | 2.4000023459279602 | 2.3000082025538129 | 2.2000018654066062 | 2.0999912335446838 | 2.0074490444977489 | 1.9404816204440281 | 1.8313385741140977 | 1.7064019721397619 | 1.6094876898532873 | 1.506829617234553 | 1.4053634378769695 | 1.308457694189304 | 1.2057996215705744 | 1.1043505192427383 | 1.0002604985473884 | 6 | 5.9030899869919438 | 5.8239109124112867 | 5.6989700043360099 | 5.6020599913279625 | 5.4881133818002974 | 5.3979400086720375 | 5.301029995663991 | 5.204119982655925 | 5.0969100130080545 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521985 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.301029995663991 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.9004927554293745 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463871 | 3.4001233554370853 | 3.3001605369513567 | 3.2001388538573878 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.800036241827986 | 2.7000023971422542 | 2.6000249038929892 | 2.5000030680516976 | 2.4000023459279602 | 2.3000082025538129 | 2.2000018654066062 | 2.0999912335446838 | 2.0074490444977489 | 1.9404816204440281 | 1.8313385741140977 | 1.7064019721397619 | 1.6094876898532873 | 1.506829617234553 | 1.4053634378769695 | 1.308457694189304 | 1.2057996215705744 | 1.1043505192427383 | 1.0002604985473884 | 6 | 5.9030899869919438 | 5.8239109124112867 | 5.6989700043360099 | 5.6020599913279625 | 5.4881133818002974 | 5.3979400086720375 | 5.301029995663991 | 5.204119982655925 | 5.0969100130080545 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521985 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.301029995663991 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.9004927554293745 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463871 | 3.4001233554370853 | 3.3001605369513567 | 3.2001388538573878 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.800036241827986
5 h | #N/A | #N/A | #N/A | #N/A | #N/A | -3.1102 | -3.1009 | -3.0405 | -2.9908 | -2.918 | -2.8477 | -2.774 | -2.6924 | -2.6069 | -2.5169 | -2.4257 | -2.3303 | -2.2342 | -2.1354 | -2.0347 | -1.929 | -1.825 | -1.7117 | -1.592 | -1.4692 | -1.342 | -1.2159 | -1.0976 | -0.99409 | -0.91022 | -0.85107 | -0.81805 | -0.81083 | -0.82709 | -0.86333 | -0.91551 | -0.97972 | -1.0525 | -1.1308 | -1.2118 | -1.2876 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
4 h | #N/A | #N/A | #N/A | #N/A | #N/A | -3.1102 | -3.1009 | -3.0405 | -2.9908 | -2.918 | -2.8477 | -2.774 | -2.6924 | -2.6069 | -2.5169 | -2.4257 | -2.3303 | -2.2342 | -2.1354 | -2.0347 | -1.929 | -1.825 | -1.7117 | -1.592 | -1.4692 | -1.342 | -1.2159 | -1.0976 | -0.99409 | -0.91022 | -0.85107 | -0.81805 | -0.81083 | -0.82709 | -0.86333 | -0.91551 | -0.97972 | -1.0525 | -1.1308 | -1.2118 | -1.2876 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
3 h | #N/A | #N/A | #N/A | #N/A | #N/A | -3.1019 | -3.09 | -3.0296 | -2.9785 | -2.9049 | -2.8336 | -2.7596 | -2.6777 | -2.5918 | -2.5014 | -2.4099 | -2.3141 | -2.2177 | -2.1185 | -2.0174 | -1.9111 | -1.8066 | -1.6925 | -1.5721 | -1.4488 | -1.3216 | -1.1964 | -1.0799 | -0.9791 | -0.89875 | -0.84361 | -0.8148 | -0.81166 | -0.83165 | -0.87093 | -0.9256 | -0.99166 | -1.0657 | -1.1447 | -1.226 | -1.3023 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
2 h | #N/A | #N/A | #N/A | #N/A | #N/A | -3.0812 | -3.0667 | -3.0052 | -2.9526 | -2.8779 | -2.8059 | -2.7305 | -2.648 | -2.5614 | -2.4701 | -2.3779 | -2.2815 | -2.1845 | -2.0845 | -1.9825 | -1.8794 | -1.7694 | -1.6539 | -1.5321 | -1.408 | -1.2813 | -1.158 | -1.046 | -0.9506 | -0.87763 | -0.83086 | -0.81055 | -0.81552 | -0.84263 | -0.88786 | -0.94711 | -1.0165 | -1.093 | -1.1728 | -1.2545 | -1.3263 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
1.5 h | #N/A | #N/A | #N/A | #N/A | #N/A | -3.0566 | -3.0404 | -2.978 | -2.923 | -2.8467 | -2.7733 | -2.6971 | -2.613 | -2.5255 | -2.4333 | -2.3404 | -2.2433 | -2.1455 | -2.0447 | -1.9417 | -1.8372 | -1.7257 | -1.6085 | -1.4853 | -1.3606 | -1.2347 | -1.1146 | -1.008 | -0.92019 | -0.85641 | -0.8196 | -0.80928 | -0.8234 | -0.85836 | -0.90989 | -0.97384 | -1.0466 | -1.1245 | -1.2054 | -1.2868 | -1.3604 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
1 h | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | -2.8376 | -2.7615 | -2.6912 | -2.601 | -2.5165 | -2.4303 | -2.3378 | -2.2435 | -2.144 | -2.0446 | -1.9448 | -1.8395 | -1.7293 | -1.6149 | -1.4952 | -1.3721 | -1.2479 | -1.127 | -1.0188 | -0.92761 | -0.85977 | -0.81926 | -0.80592 | -0.8179 | -0.85114 | -0.90078 | -0.96208 | -1.0306 | -1.1023 | -1.1736 | -1.2409 | -1.3009 | -1.3506 | -1.3857 | -1.4052 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
0.5 h | #N/A | #N/A | -1.5019 | -1.3853 | -1.2877 | -1.1767 | -1.0986 | -1.0149 | -0.93803 | -0.8642 | -0.81165 | -0.77674 | -0.76175 | -0.76911 | -0.7986 | -0.84516 | -0.90555 | -0.97431 | -1.049 | -1.1259 | -1.2043 | -1.2803 | -1.3539 | -1.4243 | -1.4894 | -1.5492 | -1.6031 | -1.6507 | -1.6919 | -1.7273 | -1.7567 | -1.7812 | -1.802 | -1.819 | -1.8348 | -1.8479 | -1.8588 | -1.8683 | -1.8781 | -1.8839 | -1.9 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
0 h | #N/A | #N/A | -1.2171 | -1.1024 | -1.014 | -0.91779 | -0.85351 | -0.79541 | -0.75401 | -0.73104 | -0.73262 | -0.75468 | -0.79543 | -0.84904 | -0.91523 | -0.98583 | -1.0594 | -1.1324 | -1.2038 | -1.2712 | -1.3348 | -1.3914 | -1.4416 | -1.4851 | -1.5209 | -1.5497 | -1.5717 | -1.5874 | -1.5977 | -1.6033 | -1.6048 | -1.6029 | -1.598 | -1.5903 | -1.58 | -1.5666 | -1.5501 | -1.5315 | -1.5108 | -1.4869 | -1.4623 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
C0
C1
R1

## Slide 2
The first part of LyoSense™ comprises a modified glass freeze-drying vial, with an electrode system deposited on the external surface (Figure 3). The thermal mass of the vial-electrode-connector-system is <0.5% of the mass of the vial/product.
Figure 4 FIVE-channel impedance analyser located on the top of the freeze-drier
The second part of LyoSense™ comprises a multi-channel, high precision impedance analyzer (Figure 4).The system has 5 channels, which connect via a junction box (Figure 6) to 5 individual LyoSense™ vials positioned around the freeze-drier shelf (Figure 5)
 |  |  | 1 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 5 |  |  |  |  | 
 |  |  |  |  |  |  |  |  |  |  | 4 |  |  |  |  |  |  |  |  |  |  |  |  | 
 |  |  |  | 2 |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 3 |  |  |  | 
Frequency scans (10 Hz – 1 MHz) of the system impedance were recorded every 10 seconds during a range of freeze-drying cycles and placebo formulations (Data shown in the following figures relate to lactose 3% w/v in RO water).
The LyoSense™ measurement and control software saves the spectra from each time point in a file for further analysis by another bespoke piece of software, called LyoView™.
A typical response surface is shown in Figure 8, in which the imaginary capacitance (C″, dielectric loss) is plotted as a function of time and log frequency. C″ is characterized by a peak in the frequency spectrum which arises from the composite capacitor of the product in series with the glass vial.
Features of the response surface: Region A shows the freezing stage and Region B which shows the primary drying stage. Other regions mainly show artefacts from the measurement system and should be ignored.
The LyoView™ software analyses the amplitude (C″peak) and position on the frequency spectrum (Fpeak) and plots these as a function of time (Figure 9,10). During the freezing & annealing stages there are dramatic shifts in the frequency position and height of the peak. During primary drying, the peak frequency remains unchanged but the amplitude decreases progressively. At the end of primary drying the peak frequency increases once again.
Figure 3 Electrode system deposited onto the external surface of the sample vial
Figure 5 Disposition of individual FD vials on the shelf of the Heto8 freeze-drier.
Positions in yellow mark location of LyoSense™ test vials
(channels 1-5).
Figure 6 Junction box to connect the leads from the individual vials with the LyoSense measurement unit ()located outside of the freeze-drier)
Figure 7 Front view of the Heto 8 freeze-drier showing the test vials on the top shelf
Figure 9 Peak amplitude (A) and peak frequency (B) as a function of time for the freeze-drying of lactose 3%
Figure 8 Front view of the Heto 8 freeze-drier showing the test vials on the top shelf
Figure 10 Log peak amplitude and peak frequency as a function of time for the freeze-drying of lactose 3%
A
B
C″
(B)
Time (h)
Log frequency
Chart: Log fpeak/Hz
Chart: Set T/◦C
| 0 | 2.5 | 5 | 5.0166666666666684 | 8.0166666666666728 | 8.0333333333333332 | 23.033333333333303 | 24.033333333333303 | 25.033333333333303 | 25.366666666666667 | 28.866666666666667 | 29.2 | #N/A | #N/A | #N/A | #N/A
Set temperature (C) | 27 | -30 | -30 | -10 | -10 | -20 | -20 | 20 | 20 | 40 | 40 | 20 | 0 | 0 | 0 | 0
Chart: C" peak/pF
(A)
Chart: Log C"peak/pF
Chart: Set T/◦C
| 0 | 2.5 | 5 | 5.0166666666666684 | 8.0166666666666728 | 8.0333333333333332 | 23.033333333333303 | 24.033333333333303 | 25.033333333333303 | 25.366666666666667 | 28.866666666666667 | 29.2 | #N/A | #N/A | #N/A | #N/A
Set temperature (C) | 27 | -30 | -30 | -10 | -10 | -20 | -20 | 20 | 20 | 40 | 40 | 20 | 0 | 0 | 0 | 0
Chart: Log fpeak/Hz
Time (h)
Time (h)
(A)
(B)

## Slide 3
6. RESULTS AND DISCUSSION
Changes in phase (e.g. ice formation) and completion of freezing, the onset of micro-collapse, and the end-points of primary and secondary drying are all detected by this method through changes in the pseudo-relaxation peak of the vial-electrode assembly.
It has been demonstrated, in general, that the frequency of the peak is strongly correlated with the temperature of the product (through the interdependency of product temperature and/or phase on the electrical resistance of the product) whereas the height of the peak is dependent on the amount of ice remaining (in primary drying) and the residual surface moisture (in secondary drying). With appropriate calibration these parameters can serve as direct measures of sample moisture and temperature respectively .
As the cycle progresses the peak height decreases with a characteristic sigmoid time-dependence (Figure 9A) such that the derivative of the time-profile could be used to define the end of the primary drying (Figure 11).
Figure 11 Peak amplitude derivative throughout the primary drying phase. The maxima observed for each vial demonstrates that the ice front has receded through half the depth of the material, and at some point subsequent to the peak onset time, switching to secondary drying may be undertaken without risking collapse of the product
With the exception of the front right vial, all samples progressed to secondary drying without collapse.
ACKNOWLEDGMENTS
The Technology Strategy Board for co-funding this work through the High Value Manufacturing Collaborative R&D programme.
The sigmoid shape of the drying curve can be explained as follows:
Initially the surface of sublimation is restricted to top surface of a frozen sample, which in turn restricts the rate of drying.
As sublimation increases the upper layers of a sample become porous and sublimation penetrates deeper into sample thus increasing the area of sublimation. This reflected by the growth of the derivative.
The growth of derivative and hence the area of sublimation continues until the whole matrix of the sample from top to bottom is porous. The surface stops increasing and the derivative reach its maximum . The more porous the matrix the longer the diffusion path for water vapours. This is coupled with the fact that there is inevitably less residual ice left, as sublimation nears the end point. Both factors then contribute to a reduced rate of drying towards the end of the primary drying cycle and hence a reduction in the derivative of the peak amplitude (C″peak).
There is an additional sensitivity to the position of the ice front in relation to the position of the external electrodes. When the ice front sits either above the top of the electrode or below the bottom of the electrode, the sensitivity to changes in height of the ice layer is reduced.
7. CONCLUSIONS
In-process control may be established using the LyoSense™ system, through the definition of set-points (i.e. the product temperature during primary drying, to drive the process at a high temperature while avoiding collapse) and end points (to establish the end of primary drying and/or moisture content of the product).
8. REFERENCES
[1] T.A. Jennings. Lyophilization, 1999, InterPharm Press, Denver; p.615-617
[2] A.A. Barresi et al. “In-line control of the lyophilization process. Gentle PAT approach using software sensors” Int. J. Refrigeration 32 (2009) 1003-1014
Chart: Time/h