## Slide 1: Electrical Impedance Methods for Developing a Lyophilization Cycle
Prof. Geoff Smith and Yowwares Jeeraruangrattana
School of Pharmacy, De Montfort University, Leicester LE1 9BH
ISLFD 2019 – 9th International Symposium on Lyophilization of Pharmaceuticals
Ghent, Belgium, 2-6 September 2019

## Slide 2: Overview
Critical parameters in freezing
On-line impedance spectroscopy (TVIS)
Dielectric loss / dielectric relaxation processes (liquid to frozen)
Dielectric loss or dielectric permittivity analysis?
Dielectric permittivity spectrum: What frequency?
In-vial determination of ……
  Ice solidification rate
  Ice nucleation temperature ()
  Eutectic melting () or glass transition temperature ()

## Slide 3: Critical Parameters
Ice crystal structure (defined by freezing process and formulation)
  Dry layer resistance impacting primary drying rate
  Surface area of dry later impacting secondary drying rate
  High temperature nucleation and slow cooling favours larger crystals
  Low temperature nucleation and fast cooling favours smaller crystals
  B Scutella

## Slide 4: Through Vial Impedance Spectroscopy
Non- perturbing to packing of vials
Low thermal mass of electrodes
no interference with heat transfer & drying rates
Multichannel
Thin flexible cables (0.5 - 2 m)
Stoppering unaffected
Single Vial PAT
Non-sample invasive
no impact on ice nucleation
Temperature calibration
using nearest neighbour vial(s)

## Slide 5: Dielectric Loss/Relaxation Mechanisms
Through Vial Impedance Spectroscopy (TVIS)
Dielectric Loss/Relaxation Mechanisms

## Slide 6: Dielectric Loss Mechanisms
The polarization of the water dipole in liquid water at 20 ˚C, with a dielectric loss peak frequency of ~ 18 GHz
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0176 | 1.024 | 1.0293 | 1.0371 | 1.0465 | 1.0575 | 1.0742 | 1.0971 | 1.131 | 1.18 | 1.2538 | 1.3532 | 1.4803 | 1.6281 | 1.7827 | 1.9288 | 2.0567 | 2.1587 | 2.2374 | 2.2971 | 2.3414 | 2.3756 | 2.4038 | 2.4272 | 2.4476 | 2.4662 | 2.4852 | 2.5003 | 2.5152 | 2.5286 | 2.5419 | 2.5531 | 2.5615 | 2.5717 | 2.5802 | 2.589 | 2.5992 | 2.6061 | 2.6201 | 2.625 | 2.6478 | 2.6645 | 2.686 | 2.708 | 2.7359 | 2.7645 | 2.7969
Time 5.03 h, -20 celcius | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.083003 | 0.084127 | 0.084469 | 0.085543 | 0.086743 | 0.087572 | 0.088047 | 0.088435 | 0.088561 | 0.088505 | 0.088289 | 0.08811 | 0.08822 | 0.08935 | 0.092855 | 0.10126 | 0.11875 | 0.15045 | 0.2025 | 0.28088 | 0.38519 | 0.51085 | 0.64628 | 0.77735 | 0.89303 | 0.98847 | 1.0619 | 1.1159 | 1.155 | 1.183 | 1.203 | 1.2167 | 1.2267 | 1.2373 | 1.2444 | 1.2521 | 1.2603 | 1.2702 | 1.2796 | 1.2944 | 1.3089 | 1.3285 | 1.3531 | 1.3812 | 1.4143 | 1.4539 | 1.499
Time 8.5 h, -40 celcius | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.08324 | 0.083622 | 0.083847 | 0.084834 | 0.086128 | 0.086714 | 0.087381 | 0.087886 | 0.088558 | 0.089009 | 0.089429 | 0.089952 | 0.090576 | 0.09152 | 0.092833 | 0.094685 | 0.097293 | 0.10083 | 0.10562 | 0.11226 | 0.12153 | 0.13502 | 0.15512 | 0.18502 | 0.22832 | 0.28819 | 0.36571 | 0.45912 | 0.56375 | 0.67242 | 0.77798 | 0.87404 | 0.95702 | 1.0284 | 1.0845 | 1.1268 | 1.1614 | 1.1873 | 1.2052 | 1.2167 | 1.2287 | 1.2359 | 1.2394 | 1.2468 | 1.2512 | 1.2552 | 1.2597
Std. : Water (1:0)_ 5 g_17C | 2.3304 | 2.4387 | 2.4947 | 2.5334 | 2.5734 | 2.595 | 2.6101 | 2.6237 | 2.6363 | 2.6458 | 2.6525 | 2.6605 | 2.6675 | 2.6754 | 2.6841 | 2.6936 | 2.7028 | 2.7121 | 2.7221 | 2.732 | 2.7416 | 2.7519 | 2.7627 | 2.773 | 2.7842 | 2.7949 | 2.8043 | 2.8148 | 2.8246 | 2.8354 | 2.8466 | 2.8598 | 2.8739 | 2.8899 | 2.908 | 2.9285 | 2.9493 | 2.972 | 2.9971 | 3.0212 | 3.0504 | 3.076 | 3.1069 | 3.1381 | 3.1714 | 3.2072 | 3.2459 | 3.2864
Std. : Water (1:3)_5 g_22.3C | 1.4671 | 1.6855 | 1.744 | 1.8824 | 1.9929 | 2.1292 | 2.3085 | 2.4363 | 2.5381 | 2.631 | 2.6959 | 2.7426 | 2.7767 | 2.8037 | 2.8219 | 2.8373 | 2.8512 | 2.8646 | 2.8761 | 2.8875 | 2.8997 | 2.9118 | 2.9235 | 2.9355 | 2.9478 | 2.9606 | 2.9734 | 2.9866 | 2.9991 | 3.0123 | 3.0249 | 3.039 | 3.0531 | 3.0685 | 3.0846 | 3.1036 | 3.1223 | 3.1442 | 3.1674 | 3.1919 | 3.2199 | 3.2461 | 3.2797 | 3.3073 | 3.3426 | 3.3775 | 3.4158 | 3.4569 | 3.5017 | 3.5489
Fitting-Time 0.23 h (20.3 C) Liquid | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.4119 | 2.4119 | 2.4119 | 2.4119 | 2.4118 | 2.4117 | 2.4115 | 2.4112 | 2.4108 | 2.4101 | 2.409 | 2.4073 | 2.4046 | 2.4004 | 2.3937 | 2.3832 | 2.367 | 2.342 | 2.3043 | 2.2484 | 2.1697 | 2.0642 | 1.9301 | 1.7781 | 1.6192 | 1.4715 | 1.3445 | 1.2461 | 1.1731 | 1.1243 | 1.0907 | 1.0692 | 1.0535 | 1.0451 | 1.0395 | 1.035 | 1.0329 | 1.0318 | 1.0305 | 1.0302 | 1.0299 | 1.0297
Fitting-Time 5.03 h (-20 C) Frozen | 1.2389 | 1.2389 | 1.2389 | 1.2389 | 1.2388 | 1.2387 | 1.2386 | 1.2384 | 1.2381 | 1.2377 | 1.2369 | 1.2358 | 1.2339 | 1.231 | 1.2264 | 1.2192 | 1.208 | 1.1906 | 1.1642 | 1.1248 | 1.0679 | 0.98926 | 0.88724 | 0.76529 | 0.63274 | 0.50242 | 0.38716 | 0.29408 | 0.22349 | 0.17372 | 0.13979 | 0.11692 | 0.10222 | 0.092682 | 0.086624 | 0.08272 | 0.080267 | 0.078693 | 0.077739 | 0.077121 | 0.076743 | 0.076474 | 0.076332 | 0.076237 | 0.076164 | 0.07613 | 0.076111 | 0.076089 | 0.076084 | 0.076079 | 0.076076
18 GHz
TVIS range
~ 80
~ 3
https://en.wikipedia.org/wiki/Permittivity
Real part Capacitance
Imaginary part Capacitance
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 0.021874 | 0.027746 | 0.035469 | 0.045629 | 0.058992 | 0.069416 | 0.087012 | 0.11672 | 0.14702 | 0.18167 | 0.23092 | 0.28348 | 0.3471 | 0.41787 | 0.49639 | 0.57052 | 0.6315 | 0.66609 | 0.66868 | 0.64015 | 0.58783 | 0.52385 | 0.45806 | 0.3962 | 0.3406 | 0.2936 | 0.25428 | 0.22269 | 0.19622 | 0.17597 | 0.15804 | 0.14279 | 0.13021 | 0.12053 | 0.11142 | 0.10593 | 0.099251 | 0.10112 | 0.096409 | 0.099903 | 0.099379 | 0.11049 | 0.12227 | 0.13167 | 0.14829 | 0.16415 | 0.18118 | 0.20571 | 0.22768 | 0.25157 | 0.27443
Time 5.03 h, -20 celcius | 0.0046903 | 0.0060276 | 0.0048957 | 0.0052689 | 0.0080706 | 0.0070995 | 0.0085621 | 0.011385 | 0.014441 | 0.017939 | 0.023068 | 0.028754 | 0.036271 | 0.045614 | 0.058069 | 0.073874 | 0.094576 | 0.12111 | 0.15588 | 0.20027 | 0.2566 | 0.32289 | 0.39693 | 0.47149 | 0.53464 | 0.57592 | 0.58749 | 0.56901 | 0.52696 | 0.47094 | 0.40966 | 0.35018 | 0.29633 | 0.24993 | 0.21217 | 0.18272 | 0.16189 | 0.14492 | 0.13366 | 0.12856 | 0.12752 | 0.12988 | 0.13813 | 0.14441 | 0.16577 | 0.18416 | 0.20379 | 0.22754 | 0.25211 | 0.27626 | 0.29768
Time 8.5 h, -40 celcius | 0.0037257 | 0.0037257 | 0.0022099 | 0.0055149 | 0.0027352 | 0.0046061 | 0.0028107 | 0.0031095 | 0.0040131 | 0.0050574 | 0.0061923 | 0.007861 | 0.0096387 | 0.01196 | 0.014763 | 0.018394 | 0.022849 | 0.028488 | 0.035374 | 0.043956 | 0.054453 | 0.067485 | 0.083211 | 0.10248 | 0.12627 | 0.15518 | 0.19059 | 0.23316 | 0.28257 | 0.33712 | 0.39303 | 0.44437 | 0.48479 | 0.5087 | 0.51347 | 0.50011 | 0.47159 | 0.43276 | 0.38638 | 0.3379 | 0.29199 | 0.24717 | 0.21022 | 0.17522 | 0.14921 | 0.12645 | 0.10764 | 0.094062 | 0.085294 | 0.077832 | 0.076166 | 0.074374
Std. : Water (1:0)_ 5 g_17C | 0.40745 | 0.37758 | 0.33167 | 0.27797 | 0.22443 | 0.19417 | 0.16169 | 0.13849 | 0.12143 | 0.10697 | 0.096405 | 0.087034 | 0.081871 | 0.078656 | 0.076272 | 0.075226 | 0.07448 | 0.073601 | 0.073296 | 0.073701 | 0.074654 | 0.075483 | 0.075353 | 0.07631 | 0.0769 | 0.077719 | 0.079173 | 0.081297 | 0.084848 | 0.090883 | 0.098449 | 0.10715 | 0.11892 | 0.13057 | 0.14131 | 0.15224 | 0.161 | 0.1718 | 0.18315 | 0.19373 | 0.20789 | 0.22266 | 0.23803 | 0.25543 | 0.2743 | 0.29629 | 0.32039 | 0.34659
Std. : Water (1:3)_5 g_22.3C | 0.35022 | 0.37378 | 0.41173 | 0.55328 | 0.59494 | 0.61621 | 0.60526 | 0.57334 | 0.52449 | 0.45124 | 0.3927 | 0.33463 | 0.28521 | 0.24069 | 0.2067 | 0.17971 | 0.15895 | 0.14354 | 0.13098 | 0.1209 | 0.11373 | 0.10865 | 0.10508 | 0.10263 | 0.10025 | 0.099388 | 0.099009 | 0.09943 | 0.10029 | 0.10249 | 0.10653 | 0.11105 | 0.11816 | 0.12566 | 0.13766 | 0.14944 | 0.16023 | 0.17158 | 0.18231 | 0.19465 | 0.20803 | 0.22162 | 0.23412 | 0.25639 | 0.27568 | 0.29801 | 0.32242 | 0.35086 | 0.38221 | 0.41669
Fitting-Time 0.23 h (20.3 C) Liquid | 0.00080148 | 0.001009 | 0.0012703 | 0.0015992 | 0.0020132 | 0.0025345 | 0.0031908 | 0.0040169 | 0.005057 | 0.0063663 | 0.0080146 | 0.010089 | 0.012702 | 0.01599 | 0.020128 | 0.025337 | 0.031892 | 0.040136 | 0.050507 | 0.063538 | 0.07988 | 0.10037 | 0.126 | 0.15786 | 0.1972 | 0.24538 | 0.30315 | 0.37047 | 0.44655 | 0.52567 | 0.59992 | 0.65886 | 0.68897 | 0.68388 | 0.64493 | 0.58011 | 0.50283 | 0.42213 | 0.34988 | 0.28493 | 0.23165 | 0.18163 | 0.14738 | 0.11839 | 0.089082 | 0.071374 | 0.059526 | 0.041708 | 0.035758 | 0.029805 | 0.023848
Fitting-Time 5.03 h (-20 C) Frozen | 0.0048296 | 0.0060799 | 0.0076542 | 0.0096359 | 0.01213 | 0.01527 | 0.019222 | 0.024195 | 0.030452 | 0.038322 | 0.048213 | 0.060635 | 0.076215 | 0.095709 | 0.12001 | 0.15014 | 0.18717 | 0.232 | 0.28517 | 0.34602 | 0.41191 | 0.4775 | 0.53412 | 0.57135 | 0.58091 | 0.56037 | 0.51476 | 0.45386 | 0.38691 | 0.32252 | 0.26465 | 0.2141 | 0.17243 | 0.138 | 0.11029 | 0.087693 | 0.069745 | 0.055173 | 0.044035 | 0.034967 | 0.027983 | 0.021692 | 0.017495 | 0.013998 | 0.010499 | 0.0083994 | 0.0069995 | 0.0048997 | 0.0041998 | 0.0034999 | 0.0027999
(I)
Notes: D:\PhD\Assignment\Conference ppt\Plot used in smi 2019 Conference_v1

## Slide 7: Dielectric Loss Mechanisms
Maxwell-Wagner (MW) polarization of the glass wall of the TVIS vial at +20 ˚C, with a dielectric loss peak frequency of 17.8 kHz
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0176 | 1.024 | 1.0293 | 1.0371 | 1.0465 | 1.0575 | 1.0742 | 1.0971 | 1.131 | 1.18 | 1.2538 | 1.3532 | 1.4803 | 1.6281 | 1.7827 | 1.9288 | 2.0567 | 2.1587 | 2.2374 | 2.2971 | 2.3414 | 2.3756 | 2.4038 | 2.4272 | 2.4476 | 2.4662 | 2.4852 | 2.5003 | 2.5152 | 2.5286 | 2.5419 | 2.5531 | 2.5615 | 2.5717 | 2.5802 | 2.589 | 2.5992 | 2.6061 | 2.6201 | 2.625 | 2.6478 | 2.6645 | 2.686 | 2.708 | 2.7359 | 2.7645 | 2.7969
Time 5.03 h, -20 celcius | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.083003 | 0.084127 | 0.084469 | 0.085543 | 0.086743 | 0.087572 | 0.088047 | 0.088435 | 0.088561 | 0.088505 | 0.088289 | 0.08811 | 0.08822 | 0.08935 | 0.092855 | 0.10126 | 0.11875 | 0.15045 | 0.2025 | 0.28088 | 0.38519 | 0.51085 | 0.64628 | 0.77735 | 0.89303 | 0.98847 | 1.0619 | 1.1159 | 1.155 | 1.183 | 1.203 | 1.2167 | 1.2267 | 1.2373 | 1.2444 | 1.2521 | 1.2603 | 1.2702 | 1.2796 | 1.2944 | 1.3089 | 1.3285 | 1.3531 | 1.3812 | 1.4143 | 1.4539 | 1.499
Time 8.5 h, -40 celcius | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.08324 | 0.083622 | 0.083847 | 0.084834 | 0.086128 | 0.086714 | 0.087381 | 0.087886 | 0.088558 | 0.089009 | 0.089429 | 0.089952 | 0.090576 | 0.09152 | 0.092833 | 0.094685 | 0.097293 | 0.10083 | 0.10562 | 0.11226 | 0.12153 | 0.13502 | 0.15512 | 0.18502 | 0.22832 | 0.28819 | 0.36571 | 0.45912 | 0.56375 | 0.67242 | 0.77798 | 0.87404 | 0.95702 | 1.0284 | 1.0845 | 1.1268 | 1.1614 | 1.1873 | 1.2052 | 1.2167 | 1.2287 | 1.2359 | 1.2394 | 1.2468 | 1.2512 | 1.2552 | 1.2597
Std. : Water (1:0)_ 5 g_17C | 2.3304 | 2.4387 | 2.4947 | 2.5334 | 2.5734 | 2.595 | 2.6101 | 2.6237 | 2.6363 | 2.6458 | 2.6525 | 2.6605 | 2.6675 | 2.6754 | 2.6841 | 2.6936 | 2.7028 | 2.7121 | 2.7221 | 2.732 | 2.7416 | 2.7519 | 2.7627 | 2.773 | 2.7842 | 2.7949 | 2.8043 | 2.8148 | 2.8246 | 2.8354 | 2.8466 | 2.8598 | 2.8739 | 2.8899 | 2.908 | 2.9285 | 2.9493 | 2.972 | 2.9971 | 3.0212 | 3.0504 | 3.076 | 3.1069 | 3.1381 | 3.1714 | 3.2072 | 3.2459 | 3.2864
Std. : Water (1:3)_5 g_22.3C | 1.4671 | 1.6855 | 1.744 | 1.8824 | 1.9929 | 2.1292 | 2.3085 | 2.4363 | 2.5381 | 2.631 | 2.6959 | 2.7426 | 2.7767 | 2.8037 | 2.8219 | 2.8373 | 2.8512 | 2.8646 | 2.8761 | 2.8875 | 2.8997 | 2.9118 | 2.9235 | 2.9355 | 2.9478 | 2.9606 | 2.9734 | 2.9866 | 2.9991 | 3.0123 | 3.0249 | 3.039 | 3.0531 | 3.0685 | 3.0846 | 3.1036 | 3.1223 | 3.1442 | 3.1674 | 3.1919 | 3.2199 | 3.2461 | 3.2797 | 3.3073 | 3.3426 | 3.3775 | 3.4158 | 3.4569 | 3.5017 | 3.5489
Fitting-Time 0.23 h (20.3 C) Liquid | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.4119 | 2.4119 | 2.4119 | 2.4119 | 2.4118 | 2.4117 | 2.4115 | 2.4112 | 2.4108 | 2.4101 | 2.409 | 2.4073 | 2.4046 | 2.4004 | 2.3937 | 2.3832 | 2.367 | 2.342 | 2.3043 | 2.2484 | 2.1697 | 2.0642 | 1.9301 | 1.7781 | 1.6192 | 1.4715 | 1.3445 | 1.2461 | 1.1731 | 1.1243 | 1.0907 | 1.0692 | 1.0535 | 1.0451 | 1.0395 | 1.035 | 1.0329 | 1.0318 | 1.0305 | 1.0302 | 1.0299 | 1.0297
Fitting-Time 5.03 h (-20 C) Frozen | 1.2389 | 1.2389 | 1.2389 | 1.2389 | 1.2388 | 1.2387 | 1.2386 | 1.2384 | 1.2381 | 1.2377 | 1.2369 | 1.2358 | 1.2339 | 1.231 | 1.2264 | 1.2192 | 1.208 | 1.1906 | 1.1642 | 1.1248 | 1.0679 | 0.98926 | 0.88724 | 0.76529 | 0.63274 | 0.50242 | 0.38716 | 0.29408 | 0.22349 | 0.17372 | 0.13979 | 0.11692 | 0.10222 | 0.092682 | 0.086624 | 0.08272 | 0.080267 | 0.078693 | 0.077739 | 0.077121 | 0.076743 | 0.076474 | 0.076332 | 0.076237 | 0.076164 | 0.07613 | 0.076111 | 0.076089 | 0.076084 | 0.076079 | 0.076076
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 0.021874 | 0.027746 | 0.035469 | 0.045629 | 0.058992 | 0.069416 | 0.087012 | 0.11672 | 0.14702 | 0.18167 | 0.23092 | 0.28348 | 0.3471 | 0.41787 | 0.49639 | 0.57052 | 0.6315 | 0.66609 | 0.66868 | 0.64015 | 0.58783 | 0.52385 | 0.45806 | 0.3962 | 0.3406 | 0.2936 | 0.25428 | 0.22269 | 0.19622 | 0.17597 | 0.15804 | 0.14279 | 0.13021 | 0.12053 | 0.11142 | 0.10593 | 0.099251 | 0.10112 | 0.096409 | 0.099903 | 0.099379 | 0.11049 | 0.12227 | 0.13167 | 0.14829 | 0.16415 | 0.18118 | 0.20571 | 0.22768 | 0.25157 | 0.27443
Time 5.03 h, -20 celcius | 0.0046903 | 0.0060276 | 0.0048957 | 0.0052689 | 0.0080706 | 0.0070995 | 0.0085621 | 0.011385 | 0.014441 | 0.017939 | 0.023068 | 0.028754 | 0.036271 | 0.045614 | 0.058069 | 0.073874 | 0.094576 | 0.12111 | 0.15588 | 0.20027 | 0.2566 | 0.32289 | 0.39693 | 0.47149 | 0.53464 | 0.57592 | 0.58749 | 0.56901 | 0.52696 | 0.47094 | 0.40966 | 0.35018 | 0.29633 | 0.24993 | 0.21217 | 0.18272 | 0.16189 | 0.14492 | 0.13366 | 0.12856 | 0.12752 | 0.12988 | 0.13813 | 0.14441 | 0.16577 | 0.18416 | 0.20379 | 0.22754 | 0.25211 | 0.27626 | 0.29768
Time 8.5 h, -40 celcius | 0.0037257 | 0.0037257 | 0.0022099 | 0.0055149 | 0.0027352 | 0.0046061 | 0.0028107 | 0.0031095 | 0.0040131 | 0.0050574 | 0.0061923 | 0.007861 | 0.0096387 | 0.01196 | 0.014763 | 0.018394 | 0.022849 | 0.028488 | 0.035374 | 0.043956 | 0.054453 | 0.067485 | 0.083211 | 0.10248 | 0.12627 | 0.15518 | 0.19059 | 0.23316 | 0.28257 | 0.33712 | 0.39303 | 0.44437 | 0.48479 | 0.5087 | 0.51347 | 0.50011 | 0.47159 | 0.43276 | 0.38638 | 0.3379 | 0.29199 | 0.24717 | 0.21022 | 0.17522 | 0.14921 | 0.12645 | 0.10764 | 0.094062 | 0.085294 | 0.077832 | 0.076166 | 0.074374
Std. : Water (1:0)_ 5 g_17C | 0.40745 | 0.37758 | 0.33167 | 0.27797 | 0.22443 | 0.19417 | 0.16169 | 0.13849 | 0.12143 | 0.10697 | 0.096405 | 0.087034 | 0.081871 | 0.078656 | 0.076272 | 0.075226 | 0.07448 | 0.073601 | 0.073296 | 0.073701 | 0.074654 | 0.075483 | 0.075353 | 0.07631 | 0.0769 | 0.077719 | 0.079173 | 0.081297 | 0.084848 | 0.090883 | 0.098449 | 0.10715 | 0.11892 | 0.13057 | 0.14131 | 0.15224 | 0.161 | 0.1718 | 0.18315 | 0.19373 | 0.20789 | 0.22266 | 0.23803 | 0.25543 | 0.2743 | 0.29629 | 0.32039 | 0.34659
Std. : Water (1:3)_5 g_22.3C | 0.35022 | 0.37378 | 0.41173 | 0.55328 | 0.59494 | 0.61621 | 0.60526 | 0.57334 | 0.52449 | 0.45124 | 0.3927 | 0.33463 | 0.28521 | 0.24069 | 0.2067 | 0.17971 | 0.15895 | 0.14354 | 0.13098 | 0.1209 | 0.11373 | 0.10865 | 0.10508 | 0.10263 | 0.10025 | 0.099388 | 0.099009 | 0.09943 | 0.10029 | 0.10249 | 0.10653 | 0.11105 | 0.11816 | 0.12566 | 0.13766 | 0.14944 | 0.16023 | 0.17158 | 0.18231 | 0.19465 | 0.20803 | 0.22162 | 0.23412 | 0.25639 | 0.27568 | 0.29801 | 0.32242 | 0.35086 | 0.38221 | 0.41669
Fitting-Time 0.23 h (20.3 C) Liquid | 0.00080148 | 0.001009 | 0.0012703 | 0.0015992 | 0.0020132 | 0.0025345 | 0.0031908 | 0.0040169 | 0.005057 | 0.0063663 | 0.0080146 | 0.010089 | 0.012702 | 0.01599 | 0.020128 | 0.025337 | 0.031892 | 0.040136 | 0.050507 | 0.063538 | 0.07988 | 0.10037 | 0.126 | 0.15786 | 0.1972 | 0.24538 | 0.30315 | 0.37047 | 0.44655 | 0.52567 | 0.59992 | 0.65886 | 0.68897 | 0.68388 | 0.64493 | 0.58011 | 0.50283 | 0.42213 | 0.34988 | 0.28493 | 0.23165 | 0.18163 | 0.14738 | 0.11839 | 0.089082 | 0.071374 | 0.059526 | 0.041708 | 0.035758 | 0.029805 | 0.023848
Fitting-Time 5.03 h (-20 C) Frozen | 0.0048296 | 0.0060799 | 0.0076542 | 0.0096359 | 0.01213 | 0.01527 | 0.019222 | 0.024195 | 0.030452 | 0.038322 | 0.048213 | 0.060635 | 0.076215 | 0.095709 | 0.12001 | 0.15014 | 0.18717 | 0.232 | 0.28517 | 0.34602 | 0.41191 | 0.4775 | 0.53412 | 0.57135 | 0.58091 | 0.56037 | 0.51476 | 0.45386 | 0.38691 | 0.32252 | 0.26465 | 0.2141 | 0.17243 | 0.138 | 0.11029 | 0.087693 | 0.069745 | 0.055173 | 0.044035 | 0.034967 | 0.027983 | 0.021692 | 0.017495 | 0.013998 | 0.010499 | 0.0083994 | 0.0069995 | 0.0048997 | 0.0041998 | 0.0034999 | 0.0027999
18 GHz
20.3 °C
18 kHz
Real part Capacitance
Imaginary part Capacitance
Measurement vial
+
-
-
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
Water
Glass wall
Electronic polarization
Atomic polarization
o
H
H
δ-
δ+
δ+
Dipolar polarization
Grotthus mechanism
o
H
H
δ-
δ+
δ+
o
H
H
δ-
δ+
δ+
H
o
H
H
o
H
H
o
H
H+
Ionic diffusion
H+
HCO3−
Na+
Cl−
H+
H+
OH-
OH-
H+
(I)
(II)
20.3 °C
Notes: D:\PhD\Assignment\Conference ppt\Plot used in smi 2019 Conference_v1

## Slide 8: Dielectric Loss Mechanisms
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0176 | 1.024 | 1.0293 | 1.0371 | 1.0465 | 1.0575 | 1.0742 | 1.0971 | 1.131 | 1.18 | 1.2538 | 1.3532 | 1.4803 | 1.6281 | 1.7827 | 1.9288 | 2.0567 | 2.1587 | 2.2374 | 2.2971 | 2.3414 | 2.3756 | 2.4038 | 2.4272 | 2.4476 | 2.4662 | 2.4852 | 2.5003 | 2.5152 | 2.5286 | 2.5419 | 2.5531 | 2.5615 | 2.5717 | 2.5802 | 2.589 | 2.5992 | 2.6061 | 2.6201 | 2.625 | 2.6478 | 2.6645 | 2.686 | 2.708 | 2.7359 | 2.7645 | 2.7969
Time 5.03 h, -20 celcius | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.083003 | 0.084127 | 0.084469 | 0.085543 | 0.086743 | 0.087572 | 0.088047 | 0.088435 | 0.088561 | 0.088505 | 0.088289 | 0.08811 | 0.08822 | 0.08935 | 0.092855 | 0.10126 | 0.11875 | 0.15045 | 0.2025 | 0.28088 | 0.38519 | 0.51085 | 0.64628 | 0.77735 | 0.89303 | 0.98847 | 1.0619 | 1.1159 | 1.155 | 1.183 | 1.203 | 1.2167 | 1.2267 | 1.2373 | 1.2444 | 1.2521 | 1.2603 | 1.2702 | 1.2796 | 1.2944 | 1.3089 | 1.3285 | 1.3531 | 1.3812 | 1.4143 | 1.4539 | 1.499
Time 8.5 h, -40 celcius | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.08324 | 0.083622 | 0.083847 | 0.084834 | 0.086128 | 0.086714 | 0.087381 | 0.087886 | 0.088558 | 0.089009 | 0.089429 | 0.089952 | 0.090576 | 0.09152 | 0.092833 | 0.094685 | 0.097293 | 0.10083 | 0.10562 | 0.11226 | 0.12153 | 0.13502 | 0.15512 | 0.18502 | 0.22832 | 0.28819 | 0.36571 | 0.45912 | 0.56375 | 0.67242 | 0.77798 | 0.87404 | 0.95702 | 1.0284 | 1.0845 | 1.1268 | 1.1614 | 1.1873 | 1.2052 | 1.2167 | 1.2287 | 1.2359 | 1.2394 | 1.2468 | 1.2512 | 1.2552 | 1.2597
Std. : Water (1:0)_ 5 g_17C | 2.3304 | 2.4387 | 2.4947 | 2.5334 | 2.5734 | 2.595 | 2.6101 | 2.6237 | 2.6363 | 2.6458 | 2.6525 | 2.6605 | 2.6675 | 2.6754 | 2.6841 | 2.6936 | 2.7028 | 2.7121 | 2.7221 | 2.732 | 2.7416 | 2.7519 | 2.7627 | 2.773 | 2.7842 | 2.7949 | 2.8043 | 2.8148 | 2.8246 | 2.8354 | 2.8466 | 2.8598 | 2.8739 | 2.8899 | 2.908 | 2.9285 | 2.9493 | 2.972 | 2.9971 | 3.0212 | 3.0504 | 3.076 | 3.1069 | 3.1381 | 3.1714 | 3.2072 | 3.2459 | 3.2864
Std. : Water (1:3)_5 g_22.3C | 1.4671 | 1.6855 | 1.744 | 1.8824 | 1.9929 | 2.1292 | 2.3085 | 2.4363 | 2.5381 | 2.631 | 2.6959 | 2.7426 | 2.7767 | 2.8037 | 2.8219 | 2.8373 | 2.8512 | 2.8646 | 2.8761 | 2.8875 | 2.8997 | 2.9118 | 2.9235 | 2.9355 | 2.9478 | 2.9606 | 2.9734 | 2.9866 | 2.9991 | 3.0123 | 3.0249 | 3.039 | 3.0531 | 3.0685 | 3.0846 | 3.1036 | 3.1223 | 3.1442 | 3.1674 | 3.1919 | 3.2199 | 3.2461 | 3.2797 | 3.3073 | 3.3426 | 3.3775 | 3.4158 | 3.4569 | 3.5017 | 3.5489
Fitting-Time 0.23 h (20.3 C) Liquid | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.4119 | 2.4119 | 2.4119 | 2.4119 | 2.4118 | 2.4117 | 2.4115 | 2.4112 | 2.4108 | 2.4101 | 2.409 | 2.4073 | 2.4046 | 2.4004 | 2.3937 | 2.3832 | 2.367 | 2.342 | 2.3043 | 2.2484 | 2.1697 | 2.0642 | 1.9301 | 1.7781 | 1.6192 | 1.4715 | 1.3445 | 1.2461 | 1.1731 | 1.1243 | 1.0907 | 1.0692 | 1.0535 | 1.0451 | 1.0395 | 1.035 | 1.0329 | 1.0318 | 1.0305 | 1.0302 | 1.0299 | 1.0297
Fitting-Time 5.03 h (-20 C) Frozen | 1.2389 | 1.2389 | 1.2389 | 1.2389 | 1.2388 | 1.2387 | 1.2386 | 1.2384 | 1.2381 | 1.2377 | 1.2369 | 1.2358 | 1.2339 | 1.231 | 1.2264 | 1.2192 | 1.208 | 1.1906 | 1.1642 | 1.1248 | 1.0679 | 0.98926 | 0.88724 | 0.76529 | 0.63274 | 0.50242 | 0.38716 | 0.29408 | 0.22349 | 0.17372 | 0.13979 | 0.11692 | 0.10222 | 0.092682 | 0.086624 | 0.08272 | 0.080267 | 0.078693 | 0.077739 | 0.077121 | 0.076743 | 0.076474 | 0.076332 | 0.076237 | 0.076164 | 0.07613 | 0.076111 | 0.076089 | 0.076084 | 0.076079 | 0.076076
Maxwell-Wagner (MW) polarization of the glass wall of the TVIS vial at +20 ˚C, with a dielectric loss peak frequency of 17.8 kHz
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 0.021874 | 0.027746 | 0.035469 | 0.045629 | 0.058992 | 0.069416 | 0.087012 | 0.11672 | 0.14702 | 0.18167 | 0.23092 | 0.28348 | 0.3471 | 0.41787 | 0.49639 | 0.57052 | 0.6315 | 0.66609 | 0.66868 | 0.64015 | 0.58783 | 0.52385 | 0.45806 | 0.3962 | 0.3406 | 0.2936 | 0.25428 | 0.22269 | 0.19622 | 0.17597 | 0.15804 | 0.14279 | 0.13021 | 0.12053 | 0.11142 | 0.10593 | 0.099251 | 0.10112 | 0.096409 | 0.099903 | 0.099379 | 0.11049 | 0.12227 | 0.13167 | 0.14829 | 0.16415 | 0.18118 | 0.20571 | 0.22768 | 0.25157 | 0.27443
Time 5.03 h, -20 celcius | 0.0046903 | 0.0060276 | 0.0048957 | 0.0052689 | 0.0080706 | 0.0070995 | 0.0085621 | 0.011385 | 0.014441 | 0.017939 | 0.023068 | 0.028754 | 0.036271 | 0.045614 | 0.058069 | 0.073874 | 0.094576 | 0.12111 | 0.15588 | 0.20027 | 0.2566 | 0.32289 | 0.39693 | 0.47149 | 0.53464 | 0.57592 | 0.58749 | 0.56901 | 0.52696 | 0.47094 | 0.40966 | 0.35018 | 0.29633 | 0.24993 | 0.21217 | 0.18272 | 0.16189 | 0.14492 | 0.13366 | 0.12856 | 0.12752 | 0.12988 | 0.13813 | 0.14441 | 0.16577 | 0.18416 | 0.20379 | 0.22754 | 0.25211 | 0.27626 | 0.29768
Time 8.5 h, -40 celcius | 0.0037257 | 0.0037257 | 0.0022099 | 0.0055149 | 0.0027352 | 0.0046061 | 0.0028107 | 0.0031095 | 0.0040131 | 0.0050574 | 0.0061923 | 0.007861 | 0.0096387 | 0.01196 | 0.014763 | 0.018394 | 0.022849 | 0.028488 | 0.035374 | 0.043956 | 0.054453 | 0.067485 | 0.083211 | 0.10248 | 0.12627 | 0.15518 | 0.19059 | 0.23316 | 0.28257 | 0.33712 | 0.39303 | 0.44437 | 0.48479 | 0.5087 | 0.51347 | 0.50011 | 0.47159 | 0.43276 | 0.38638 | 0.3379 | 0.29199 | 0.24717 | 0.21022 | 0.17522 | 0.14921 | 0.12645 | 0.10764 | 0.094062 | 0.085294 | 0.077832 | 0.076166 | 0.074374
Std. : Water (1:0)_ 5 g_17C | 0.40745 | 0.37758 | 0.33167 | 0.27797 | 0.22443 | 0.19417 | 0.16169 | 0.13849 | 0.12143 | 0.10697 | 0.096405 | 0.087034 | 0.081871 | 0.078656 | 0.076272 | 0.075226 | 0.07448 | 0.073601 | 0.073296 | 0.073701 | 0.074654 | 0.075483 | 0.075353 | 0.07631 | 0.0769 | 0.077719 | 0.079173 | 0.081297 | 0.084848 | 0.090883 | 0.098449 | 0.10715 | 0.11892 | 0.13057 | 0.14131 | 0.15224 | 0.161 | 0.1718 | 0.18315 | 0.19373 | 0.20789 | 0.22266 | 0.23803 | 0.25543 | 0.2743 | 0.29629 | 0.32039 | 0.34659
Std. : Water (1:3)_5 g_22.3C | 0.35022 | 0.37378 | 0.41173 | 0.55328 | 0.59494 | 0.61621 | 0.60526 | 0.57334 | 0.52449 | 0.45124 | 0.3927 | 0.33463 | 0.28521 | 0.24069 | 0.2067 | 0.17971 | 0.15895 | 0.14354 | 0.13098 | 0.1209 | 0.11373 | 0.10865 | 0.10508 | 0.10263 | 0.10025 | 0.099388 | 0.099009 | 0.09943 | 0.10029 | 0.10249 | 0.10653 | 0.11105 | 0.11816 | 0.12566 | 0.13766 | 0.14944 | 0.16023 | 0.17158 | 0.18231 | 0.19465 | 0.20803 | 0.22162 | 0.23412 | 0.25639 | 0.27568 | 0.29801 | 0.32242 | 0.35086 | 0.38221 | 0.41669
Fitting-Time 0.23 h (20.3 C) Liquid | 0.00080148 | 0.001009 | 0.0012703 | 0.0015992 | 0.0020132 | 0.0025345 | 0.0031908 | 0.0040169 | 0.005057 | 0.0063663 | 0.0080146 | 0.010089 | 0.012702 | 0.01599 | 0.020128 | 0.025337 | 0.031892 | 0.040136 | 0.050507 | 0.063538 | 0.07988 | 0.10037 | 0.126 | 0.15786 | 0.1972 | 0.24538 | 0.30315 | 0.37047 | 0.44655 | 0.52567 | 0.59992 | 0.65886 | 0.68897 | 0.68388 | 0.64493 | 0.58011 | 0.50283 | 0.42213 | 0.34988 | 0.28493 | 0.23165 | 0.18163 | 0.14738 | 0.11839 | 0.089082 | 0.071374 | 0.059526 | 0.041708 | 0.035758 | 0.029805 | 0.023848
Fitting-Time 5.03 h (-20 C) Frozen | 0.0048296 | 0.0060799 | 0.0076542 | 0.0096359 | 0.01213 | 0.01527 | 0.019222 | 0.024195 | 0.030452 | 0.038322 | 0.048213 | 0.060635 | 0.076215 | 0.095709 | 0.12001 | 0.15014 | 0.18717 | 0.232 | 0.28517 | 0.34602 | 0.41191 | 0.4775 | 0.53412 | 0.57135 | 0.58091 | 0.56037 | 0.51476 | 0.45386 | 0.38691 | 0.32252 | 0.26465 | 0.2141 | 0.17243 | 0.138 | 0.11029 | 0.087693 | 0.069745 | 0.055173 | 0.044035 | 0.034967 | 0.027983 | 0.021692 | 0.017495 | 0.013998 | 0.010499 | 0.0083994 | 0.0069995 | 0.0048997 | 0.0041998 | 0.0034999 | 0.0027999
(I)
18 GHz
20.3 °C
20.3 °C
18 kHz
0.0056 M KCl
(II)
0.0056 M KCl
TVIS range
Real part Capacitance
Imaginary part Capacitance
Measurement vial
+
-
-
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
Water
Glass wall
Electronic polarization
Atomic polarization
o
H
H
δ-
δ+
δ+
Dipolar polarization
Grotthus mechanism
o
H
H
δ-
δ+
δ+
o
H
H
δ-
δ+
δ+
H
o
H
H
o
H
H
o
H
H+
Ionic diffusion
H+
HCO3−
Na+
Cl−
H+
H+
OH-
H+
K+
K+
Cl−
Notes: D:\PhD\Assignment\Conference ppt\Plot used in smi 2019 Conference_v1

## Slide 9: Dielectric Loss Mechanisms
The dielectric polarization of ice at −20 ˚C, with a dielectric loss peak frequencies of 2.57 kHz
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0176 | 1.024 | 1.0293 | 1.0371 | 1.0465 | 1.0575 | 1.0742 | 1.0971 | 1.131 | 1.18 | 1.2538 | 1.3532 | 1.4803 | 1.6281 | 1.7827 | 1.9288 | 2.0567 | 2.1587 | 2.2374 | 2.2971 | 2.3414 | 2.3756 | 2.4038 | 2.4272 | 2.4476 | 2.4662 | 2.4852 | 2.5003 | 2.5152 | 2.5286 | 2.5419 | 2.5531 | 2.5615 | 2.5717 | 2.5802 | 2.589 | 2.5992 | 2.6061 | 2.6201 | 2.625 | 2.6478 | 2.6645 | 2.686 | 2.708 | 2.7359 | 2.7645 | 2.7969
Time 5.03 h, -20 celcius | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.083003 | 0.084127 | 0.084469 | 0.085543 | 0.086743 | 0.087572 | 0.088047 | 0.088435 | 0.088561 | 0.088505 | 0.088289 | 0.08811 | 0.08822 | 0.08935 | 0.092855 | 0.10126 | 0.11875 | 0.15045 | 0.2025 | 0.28088 | 0.38519 | 0.51085 | 0.64628 | 0.77735 | 0.89303 | 0.98847 | 1.0619 | 1.1159 | 1.155 | 1.183 | 1.203 | 1.2167 | 1.2267 | 1.2373 | 1.2444 | 1.2521 | 1.2603 | 1.2702 | 1.2796 | 1.2944 | 1.3089 | 1.3285 | 1.3531 | 1.3812 | 1.4143 | 1.4539 | 1.499
Time 8.5 h, -40 celcius | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.08324 | 0.083622 | 0.083847 | 0.084834 | 0.086128 | 0.086714 | 0.087381 | 0.087886 | 0.088558 | 0.089009 | 0.089429 | 0.089952 | 0.090576 | 0.09152 | 0.092833 | 0.094685 | 0.097293 | 0.10083 | 0.10562 | 0.11226 | 0.12153 | 0.13502 | 0.15512 | 0.18502 | 0.22832 | 0.28819 | 0.36571 | 0.45912 | 0.56375 | 0.67242 | 0.77798 | 0.87404 | 0.95702 | 1.0284 | 1.0845 | 1.1268 | 1.1614 | 1.1873 | 1.2052 | 1.2167 | 1.2287 | 1.2359 | 1.2394 | 1.2468 | 1.2512 | 1.2552 | 1.2597
Std. : Water (1:0)_ 5 g_17C | 2.3304 | 2.4387 | 2.4947 | 2.5334 | 2.5734 | 2.595 | 2.6101 | 2.6237 | 2.6363 | 2.6458 | 2.6525 | 2.6605 | 2.6675 | 2.6754 | 2.6841 | 2.6936 | 2.7028 | 2.7121 | 2.7221 | 2.732 | 2.7416 | 2.7519 | 2.7627 | 2.773 | 2.7842 | 2.7949 | 2.8043 | 2.8148 | 2.8246 | 2.8354 | 2.8466 | 2.8598 | 2.8739 | 2.8899 | 2.908 | 2.9285 | 2.9493 | 2.972 | 2.9971 | 3.0212 | 3.0504 | 3.076 | 3.1069 | 3.1381 | 3.1714 | 3.2072 | 3.2459 | 3.2864
Std. : Water (1:3)_5 g_22.3C | 1.4671 | 1.6855 | 1.744 | 1.8824 | 1.9929 | 2.1292 | 2.3085 | 2.4363 | 2.5381 | 2.631 | 2.6959 | 2.7426 | 2.7767 | 2.8037 | 2.8219 | 2.8373 | 2.8512 | 2.8646 | 2.8761 | 2.8875 | 2.8997 | 2.9118 | 2.9235 | 2.9355 | 2.9478 | 2.9606 | 2.9734 | 2.9866 | 2.9991 | 3.0123 | 3.0249 | 3.039 | 3.0531 | 3.0685 | 3.0846 | 3.1036 | 3.1223 | 3.1442 | 3.1674 | 3.1919 | 3.2199 | 3.2461 | 3.2797 | 3.3073 | 3.3426 | 3.3775 | 3.4158 | 3.4569 | 3.5017 | 3.5489
Fitting-Time 0.23 h (20.3 C) Liquid | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.4119 | 2.4119 | 2.4119 | 2.4119 | 2.4118 | 2.4117 | 2.4115 | 2.4112 | 2.4108 | 2.4101 | 2.409 | 2.4073 | 2.4046 | 2.4004 | 2.3937 | 2.3832 | 2.367 | 2.342 | 2.3043 | 2.2484 | 2.1697 | 2.0642 | 1.9301 | 1.7781 | 1.6192 | 1.4715 | 1.3445 | 1.2461 | 1.1731 | 1.1243 | 1.0907 | 1.0692 | 1.0535 | 1.0451 | 1.0395 | 1.035 | 1.0329 | 1.0318 | 1.0305 | 1.0302 | 1.0299 | 1.0297
Fitting-Time 5.03 h (-20 C) Frozen | 1.2389 | 1.2389 | 1.2389 | 1.2389 | 1.2388 | 1.2387 | 1.2386 | 1.2384 | 1.2381 | 1.2377 | 1.2369 | 1.2358 | 1.2339 | 1.231 | 1.2264 | 1.2192 | 1.208 | 1.1906 | 1.1642 | 1.1248 | 1.0679 | 0.98926 | 0.88724 | 0.76529 | 0.63274 | 0.50242 | 0.38716 | 0.29408 | 0.22349 | 0.17372 | 0.13979 | 0.11692 | 0.10222 | 0.092682 | 0.086624 | 0.08272 | 0.080267 | 0.078693 | 0.077739 | 0.077121 | 0.076743 | 0.076474 | 0.076332 | 0.076237 | 0.076164 | 0.07613 | 0.076111 | 0.076089 | 0.076084 | 0.076079 | 0.076076
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 0.021874 | 0.027746 | 0.035469 | 0.045629 | 0.058992 | 0.069416 | 0.087012 | 0.11672 | 0.14702 | 0.18167 | 0.23092 | 0.28348 | 0.3471 | 0.41787 | 0.49639 | 0.57052 | 0.6315 | 0.66609 | 0.66868 | 0.64015 | 0.58783 | 0.52385 | 0.45806 | 0.3962 | 0.3406 | 0.2936 | 0.25428 | 0.22269 | 0.19622 | 0.17597 | 0.15804 | 0.14279 | 0.13021 | 0.12053 | 0.11142 | 0.10593 | 0.099251 | 0.10112 | 0.096409 | 0.099903 | 0.099379 | 0.11049 | 0.12227 | 0.13167 | 0.14829 | 0.16415 | 0.18118 | 0.20571 | 0.22768 | 0.25157 | 0.27443
Time 5.03 h, -20 celcius | 0.0046903 | 0.0060276 | 0.0048957 | 0.0052689 | 0.0080706 | 0.0070995 | 0.0085621 | 0.011385 | 0.014441 | 0.017939 | 0.023068 | 0.028754 | 0.036271 | 0.045614 | 0.058069 | 0.073874 | 0.094576 | 0.12111 | 0.15588 | 0.20027 | 0.2566 | 0.32289 | 0.39693 | 0.47149 | 0.53464 | 0.57592 | 0.58749 | 0.56901 | 0.52696 | 0.47094 | 0.40966 | 0.35018 | 0.29633 | 0.24993 | 0.21217 | 0.18272 | 0.16189 | 0.14492 | 0.13366 | 0.12856 | 0.12752 | 0.12988 | 0.13813 | 0.14441 | 0.16577 | 0.18416 | 0.20379 | 0.22754 | 0.25211 | 0.27626 | 0.29768
Time 8.5 h, -40 celcius | 0.0037257 | 0.0037257 | 0.0022099 | 0.0055149 | 0.0027352 | 0.0046061 | 0.0028107 | 0.0031095 | 0.0040131 | 0.0050574 | 0.0061923 | 0.007861 | 0.0096387 | 0.01196 | 0.014763 | 0.018394 | 0.022849 | 0.028488 | 0.035374 | 0.043956 | 0.054453 | 0.067485 | 0.083211 | 0.10248 | 0.12627 | 0.15518 | 0.19059 | 0.23316 | 0.28257 | 0.33712 | 0.39303 | 0.44437 | 0.48479 | 0.5087 | 0.51347 | 0.50011 | 0.47159 | 0.43276 | 0.38638 | 0.3379 | 0.29199 | 0.24717 | 0.21022 | 0.17522 | 0.14921 | 0.12645 | 0.10764 | 0.094062 | 0.085294 | 0.077832 | 0.076166 | 0.074374
Std. : Water (1:0)_ 5 g_17C | 0.40745 | 0.37758 | 0.33167 | 0.27797 | 0.22443 | 0.19417 | 0.16169 | 0.13849 | 0.12143 | 0.10697 | 0.096405 | 0.087034 | 0.081871 | 0.078656 | 0.076272 | 0.075226 | 0.07448 | 0.073601 | 0.073296 | 0.073701 | 0.074654 | 0.075483 | 0.075353 | 0.07631 | 0.0769 | 0.077719 | 0.079173 | 0.081297 | 0.084848 | 0.090883 | 0.098449 | 0.10715 | 0.11892 | 0.13057 | 0.14131 | 0.15224 | 0.161 | 0.1718 | 0.18315 | 0.19373 | 0.20789 | 0.22266 | 0.23803 | 0.25543 | 0.2743 | 0.29629 | 0.32039 | 0.34659
Std. : Water (1:3)_5 g_22.3C | 0.35022 | 0.37378 | 0.41173 | 0.55328 | 0.59494 | 0.61621 | 0.60526 | 0.57334 | 0.52449 | 0.45124 | 0.3927 | 0.33463 | 0.28521 | 0.24069 | 0.2067 | 0.17971 | 0.15895 | 0.14354 | 0.13098 | 0.1209 | 0.11373 | 0.10865 | 0.10508 | 0.10263 | 0.10025 | 0.099388 | 0.099009 | 0.09943 | 0.10029 | 0.10249 | 0.10653 | 0.11105 | 0.11816 | 0.12566 | 0.13766 | 0.14944 | 0.16023 | 0.17158 | 0.18231 | 0.19465 | 0.20803 | 0.22162 | 0.23412 | 0.25639 | 0.27568 | 0.29801 | 0.32242 | 0.35086 | 0.38221 | 0.41669
Fitting-Time 0.23 h (20.3 C) Liquid | 0.00080148 | 0.001009 | 0.0012703 | 0.0015992 | 0.0020132 | 0.0025345 | 0.0031908 | 0.0040169 | 0.005057 | 0.0063663 | 0.0080146 | 0.010089 | 0.012702 | 0.01599 | 0.020128 | 0.025337 | 0.031892 | 0.040136 | 0.050507 | 0.063538 | 0.07988 | 0.10037 | 0.126 | 0.15786 | 0.1972 | 0.24538 | 0.30315 | 0.37047 | 0.44655 | 0.52567 | 0.59992 | 0.65886 | 0.68897 | 0.68388 | 0.64493 | 0.58011 | 0.50283 | 0.42213 | 0.34988 | 0.28493 | 0.23165 | 0.18163 | 0.14738 | 0.11839 | 0.089082 | 0.071374 | 0.059526 | 0.041708 | 0.035758 | 0.029805 | 0.023848
Fitting-Time 5.03 h (-20 C) Frozen | 0.0048296 | 0.0060799 | 0.0076542 | 0.0096359 | 0.01213 | 0.01527 | 0.019222 | 0.024195 | 0.030452 | 0.038322 | 0.048213 | 0.060635 | 0.076215 | 0.095709 | 0.12001 | 0.15014 | 0.18717 | 0.232 | 0.28517 | 0.34602 | 0.41191 | 0.4775 | 0.53412 | 0.57135 | 0.58091 | 0.56037 | 0.51476 | 0.45386 | 0.38691 | 0.32252 | 0.26465 | 0.2141 | 0.17243 | 0.138 | 0.11029 | 0.087693 | 0.069745 | 0.055173 | 0.044035 | 0.034967 | 0.027983 | 0.021692 | 0.017495 | 0.013998 | 0.010499 | 0.0083994 | 0.0069995 | 0.0048997 | 0.0041998 | 0.0034999 | 0.0027999
(I)
18 GHz
20.3 °C
20.3 °C
18 kHz
0.0056 M KCl
0.0056 M KCl
(II)
-20 °C
(III)
2.6 kHz
-20 °C
Real part Capacitance
Imaginary part Capacitance
Measurement vial
Co
+
-
-
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
Ice
Glass wall
Electronic polarization
Atomic polarization
L- and D- orientation defects
Notes: D:\PhD\Assignment\Conference ppt\Plot used in smi 2019 Conference_v1

## Slide 10: Dielectric Loss Mechanisms
The dielectric polarization of ice at −40 ˚C with a dielectric loss peak frequencies of 537 Hz.
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0176 | 1.024 | 1.0293 | 1.0371 | 1.0465 | 1.0575 | 1.0742 | 1.0971 | 1.131 | 1.18 | 1.2538 | 1.3532 | 1.4803 | 1.6281 | 1.7827 | 1.9288 | 2.0567 | 2.1587 | 2.2374 | 2.2971 | 2.3414 | 2.3756 | 2.4038 | 2.4272 | 2.4476 | 2.4662 | 2.4852 | 2.5003 | 2.5152 | 2.5286 | 2.5419 | 2.5531 | 2.5615 | 2.5717 | 2.5802 | 2.589 | 2.5992 | 2.6061 | 2.6201 | 2.625 | 2.6478 | 2.6645 | 2.686 | 2.708 | 2.7359 | 2.7645 | 2.7969
Time 5.03 h, -20 celcius | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.083003 | 0.084127 | 0.084469 | 0.085543 | 0.086743 | 0.087572 | 0.088047 | 0.088435 | 0.088561 | 0.088505 | 0.088289 | 0.08811 | 0.08822 | 0.08935 | 0.092855 | 0.10126 | 0.11875 | 0.15045 | 0.2025 | 0.28088 | 0.38519 | 0.51085 | 0.64628 | 0.77735 | 0.89303 | 0.98847 | 1.0619 | 1.1159 | 1.155 | 1.183 | 1.203 | 1.2167 | 1.2267 | 1.2373 | 1.2444 | 1.2521 | 1.2603 | 1.2702 | 1.2796 | 1.2944 | 1.3089 | 1.3285 | 1.3531 | 1.3812 | 1.4143 | 1.4539 | 1.499
Time 8.5 h, -40 celcius | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.08324 | 0.083622 | 0.083847 | 0.084834 | 0.086128 | 0.086714 | 0.087381 | 0.087886 | 0.088558 | 0.089009 | 0.089429 | 0.089952 | 0.090576 | 0.09152 | 0.092833 | 0.094685 | 0.097293 | 0.10083 | 0.10562 | 0.11226 | 0.12153 | 0.13502 | 0.15512 | 0.18502 | 0.22832 | 0.28819 | 0.36571 | 0.45912 | 0.56375 | 0.67242 | 0.77798 | 0.87404 | 0.95702 | 1.0284 | 1.0845 | 1.1268 | 1.1614 | 1.1873 | 1.2052 | 1.2167 | 1.2287 | 1.2359 | 1.2394 | 1.2468 | 1.2512 | 1.2552 | 1.2597
Std. : Water (1:0)_ 5 g_17C | 2.3304 | 2.4387 | 2.4947 | 2.5334 | 2.5734 | 2.595 | 2.6101 | 2.6237 | 2.6363 | 2.6458 | 2.6525 | 2.6605 | 2.6675 | 2.6754 | 2.6841 | 2.6936 | 2.7028 | 2.7121 | 2.7221 | 2.732 | 2.7416 | 2.7519 | 2.7627 | 2.773 | 2.7842 | 2.7949 | 2.8043 | 2.8148 | 2.8246 | 2.8354 | 2.8466 | 2.8598 | 2.8739 | 2.8899 | 2.908 | 2.9285 | 2.9493 | 2.972 | 2.9971 | 3.0212 | 3.0504 | 3.076 | 3.1069 | 3.1381 | 3.1714 | 3.2072 | 3.2459 | 3.2864
Std. : Water (1:3)_5 g_22.3C | 1.4671 | 1.6855 | 1.744 | 1.8824 | 1.9929 | 2.1292 | 2.3085 | 2.4363 | 2.5381 | 2.631 | 2.6959 | 2.7426 | 2.7767 | 2.8037 | 2.8219 | 2.8373 | 2.8512 | 2.8646 | 2.8761 | 2.8875 | 2.8997 | 2.9118 | 2.9235 | 2.9355 | 2.9478 | 2.9606 | 2.9734 | 2.9866 | 2.9991 | 3.0123 | 3.0249 | 3.039 | 3.0531 | 3.0685 | 3.0846 | 3.1036 | 3.1223 | 3.1442 | 3.1674 | 3.1919 | 3.2199 | 3.2461 | 3.2797 | 3.3073 | 3.3426 | 3.3775 | 3.4158 | 3.4569 | 3.5017 | 3.5489
Fitting-Time 0.23 h (20.3 C) Liquid | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.4119 | 2.4119 | 2.4119 | 2.4119 | 2.4118 | 2.4117 | 2.4115 | 2.4112 | 2.4108 | 2.4101 | 2.409 | 2.4073 | 2.4046 | 2.4004 | 2.3937 | 2.3832 | 2.367 | 2.342 | 2.3043 | 2.2484 | 2.1697 | 2.0642 | 1.9301 | 1.7781 | 1.6192 | 1.4715 | 1.3445 | 1.2461 | 1.1731 | 1.1243 | 1.0907 | 1.0692 | 1.0535 | 1.0451 | 1.0395 | 1.035 | 1.0329 | 1.0318 | 1.0305 | 1.0302 | 1.0299 | 1.0297
Fitting-Time 5.03 h (-20 C) Frozen | 1.2389 | 1.2389 | 1.2389 | 1.2389 | 1.2388 | 1.2387 | 1.2386 | 1.2384 | 1.2381 | 1.2377 | 1.2369 | 1.2358 | 1.2339 | 1.231 | 1.2264 | 1.2192 | 1.208 | 1.1906 | 1.1642 | 1.1248 | 1.0679 | 0.98926 | 0.88724 | 0.76529 | 0.63274 | 0.50242 | 0.38716 | 0.29408 | 0.22349 | 0.17372 | 0.13979 | 0.11692 | 0.10222 | 0.092682 | 0.086624 | 0.08272 | 0.080267 | 0.078693 | 0.077739 | 0.077121 | 0.076743 | 0.076474 | 0.076332 | 0.076237 | 0.076164 | 0.07613 | 0.076111 | 0.076089 | 0.076084 | 0.076079 | 0.076076
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 0.23 h, 20.3 celcius | 0.021874 | 0.027746 | 0.035469 | 0.045629 | 0.058992 | 0.069416 | 0.087012 | 0.11672 | 0.14702 | 0.18167 | 0.23092 | 0.28348 | 0.3471 | 0.41787 | 0.49639 | 0.57052 | 0.6315 | 0.66609 | 0.66868 | 0.64015 | 0.58783 | 0.52385 | 0.45806 | 0.3962 | 0.3406 | 0.2936 | 0.25428 | 0.22269 | 0.19622 | 0.17597 | 0.15804 | 0.14279 | 0.13021 | 0.12053 | 0.11142 | 0.10593 | 0.099251 | 0.10112 | 0.096409 | 0.099903 | 0.099379 | 0.11049 | 0.12227 | 0.13167 | 0.14829 | 0.16415 | 0.18118 | 0.20571 | 0.22768 | 0.25157 | 0.27443
Time 5.03 h, -20 celcius | 0.0046903 | 0.0060276 | 0.0048957 | 0.0052689 | 0.0080706 | 0.0070995 | 0.0085621 | 0.011385 | 0.014441 | 0.017939 | 0.023068 | 0.028754 | 0.036271 | 0.045614 | 0.058069 | 0.073874 | 0.094576 | 0.12111 | 0.15588 | 0.20027 | 0.2566 | 0.32289 | 0.39693 | 0.47149 | 0.53464 | 0.57592 | 0.58749 | 0.56901 | 0.52696 | 0.47094 | 0.40966 | 0.35018 | 0.29633 | 0.24993 | 0.21217 | 0.18272 | 0.16189 | 0.14492 | 0.13366 | 0.12856 | 0.12752 | 0.12988 | 0.13813 | 0.14441 | 0.16577 | 0.18416 | 0.20379 | 0.22754 | 0.25211 | 0.27626 | 0.29768
Time 8.5 h, -40 celcius | 0.0037257 | 0.0037257 | 0.0022099 | 0.0055149 | 0.0027352 | 0.0046061 | 0.0028107 | 0.0031095 | 0.0040131 | 0.0050574 | 0.0061923 | 0.007861 | 0.0096387 | 0.01196 | 0.014763 | 0.018394 | 0.022849 | 0.028488 | 0.035374 | 0.043956 | 0.054453 | 0.067485 | 0.083211 | 0.10248 | 0.12627 | 0.15518 | 0.19059 | 0.23316 | 0.28257 | 0.33712 | 0.39303 | 0.44437 | 0.48479 | 0.5087 | 0.51347 | 0.50011 | 0.47159 | 0.43276 | 0.38638 | 0.3379 | 0.29199 | 0.24717 | 0.21022 | 0.17522 | 0.14921 | 0.12645 | 0.10764 | 0.094062 | 0.085294 | 0.077832 | 0.076166 | 0.074374
Std. : Water (1:0)_ 5 g_17C | 0.40745 | 0.37758 | 0.33167 | 0.27797 | 0.22443 | 0.19417 | 0.16169 | 0.13849 | 0.12143 | 0.10697 | 0.096405 | 0.087034 | 0.081871 | 0.078656 | 0.076272 | 0.075226 | 0.07448 | 0.073601 | 0.073296 | 0.073701 | 0.074654 | 0.075483 | 0.075353 | 0.07631 | 0.0769 | 0.077719 | 0.079173 | 0.081297 | 0.084848 | 0.090883 | 0.098449 | 0.10715 | 0.11892 | 0.13057 | 0.14131 | 0.15224 | 0.161 | 0.1718 | 0.18315 | 0.19373 | 0.20789 | 0.22266 | 0.23803 | 0.25543 | 0.2743 | 0.29629 | 0.32039 | 0.34659
Std. : Water (1:3)_5 g_22.3C | 0.35022 | 0.37378 | 0.41173 | 0.55328 | 0.59494 | 0.61621 | 0.60526 | 0.57334 | 0.52449 | 0.45124 | 0.3927 | 0.33463 | 0.28521 | 0.24069 | 0.2067 | 0.17971 | 0.15895 | 0.14354 | 0.13098 | 0.1209 | 0.11373 | 0.10865 | 0.10508 | 0.10263 | 0.10025 | 0.099388 | 0.099009 | 0.09943 | 0.10029 | 0.10249 | 0.10653 | 0.11105 | 0.11816 | 0.12566 | 0.13766 | 0.14944 | 0.16023 | 0.17158 | 0.18231 | 0.19465 | 0.20803 | 0.22162 | 0.23412 | 0.25639 | 0.27568 | 0.29801 | 0.32242 | 0.35086 | 0.38221 | 0.41669
Fitting-Time 0.23 h (20.3 C) Liquid | 0.00080148 | 0.001009 | 0.0012703 | 0.0015992 | 0.0020132 | 0.0025345 | 0.0031908 | 0.0040169 | 0.005057 | 0.0063663 | 0.0080146 | 0.010089 | 0.012702 | 0.01599 | 0.020128 | 0.025337 | 0.031892 | 0.040136 | 0.050507 | 0.063538 | 0.07988 | 0.10037 | 0.126 | 0.15786 | 0.1972 | 0.24538 | 0.30315 | 0.37047 | 0.44655 | 0.52567 | 0.59992 | 0.65886 | 0.68897 | 0.68388 | 0.64493 | 0.58011 | 0.50283 | 0.42213 | 0.34988 | 0.28493 | 0.23165 | 0.18163 | 0.14738 | 0.11839 | 0.089082 | 0.071374 | 0.059526 | 0.041708 | 0.035758 | 0.029805 | 0.023848
Fitting-Time 5.03 h (-20 C) Frozen | 0.0048296 | 0.0060799 | 0.0076542 | 0.0096359 | 0.01213 | 0.01527 | 0.019222 | 0.024195 | 0.030452 | 0.038322 | 0.048213 | 0.060635 | 0.076215 | 0.095709 | 0.12001 | 0.15014 | 0.18717 | 0.232 | 0.28517 | 0.34602 | 0.41191 | 0.4775 | 0.53412 | 0.57135 | 0.58091 | 0.56037 | 0.51476 | 0.45386 | 0.38691 | 0.32252 | 0.26465 | 0.2141 | 0.17243 | 0.138 | 0.11029 | 0.087693 | 0.069745 | 0.055173 | 0.044035 | 0.034967 | 0.027983 | 0.021692 | 0.017495 | 0.013998 | 0.010499 | 0.0083994 | 0.0069995 | 0.0048997 | 0.0041998 | 0.0034999 | 0.0027999
(I)
18 GHz
20.3 °C
20.3 °C
18 kHz
(II)
-20 °C
2.6 kHz
-20 °C
-40 °C
-40 °C
537 Hz
(IV)
0.0056 M KCl
0.0056 M KCl
(III)
Real part Capacitance
Imaginary part Capacitance
Measurement vial
Co
+
-
-
+
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
+
+
+
+
+
+
+
+
+
+
-
-
-
-
-
-
-
-
-
-
Ice
Glass wall
Electronic polarization
Atomic polarization
Ionic defects
(similar to Grotthus mechanism)
Notes: D:\PhD\Assignment\Conference ppt\Plot used in smi 2019 Conference_v1

## Slide 11: Dielectric Loss Mechanisms
The polarization of the water dipole in liquid water at 20 ˚C, with a dielectric loss peak frequency of ~ 18 GHz
Maxwell-Wagner (MW) polarization of the glass wall of the TVIS vial at +20 ˚C, with a dielectric loss peak frequency of 17.8 kHz
The dielectric polarization of ice at −20 ˚C, with a dielectric loss peak frequencies of 2.57 kHz
The dielectric polarization of ice at −40 ˚C with a dielectric loss peak frequencies of 537 Hz.
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124112991 | 5.7569630370485525 | 5.6020599913279625 | 5.5228744023138026 | 5.4259741609193757 | 5.3010299956639821 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521985 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639821 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.9004927554293705 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463871 | 3.4001233554370813 | 3.3001605369513531 | 3.2001388538573816 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279803 | 2.7000023971422542 | 2.600024903892987 | 2.5000030680516936 | 2.4000023459279562 | 2.3000082025538129 | 2.2000018654066023 | 2.0999912335446842 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277406 | 1.5000030680516931 | 1.4000023459279558 | 1.3000082025538131 | 1.200001865406602 | 1.099991233544684 | 1
Time 0.23 h, 20.3 celcius | 1.0162 | 1.0162 | 1.0162 | 1.0162 | 1.0176 | 1.024 | 1.0293 | 1.0371 | 1.0465 | 1.0575 | 1.0742 | 1.0971 | 1.131 | 1.18 | 1.2538 | 1.3532 | 1.4803 | 1.6281 | 1.7827 | 1.9288 | 2.0567 | 2.1587 | 2.2374 | 2.2971 | 2.3414 | 2.3756 | 2.4038 | 2.4272 | 2.4476 | 2.4662 | 2.4852 | 2.5003 | 2.5152 | 2.5286 | 2.5419 | 2.5531 | 2.5615 | 2.5717 | 2.5802 | 2.589 | 2.5992 | 2.6061 | 2.6201 | 2.625 | 2.6478 | 2.6645 | 2.686 | 2.708 | 2.7359 | 2.7645 | 2.7969
Time 5.03 h, -20 celcius | 0.084287 | 0.084287 | 0.084287 | 0.084287 | 0.083003 | 0.084127 | 0.084469 | 0.085543 | 0.086743 | 0.087573 | 0.088047 | 0.088435 | 0.088561 | 0.088505 | 0.088289 | 0.08811 | 0.08822 | 0.08935 | 0.092855 | 0.10126 | 0.11875 | 0.15045 | 0.2025 | 0.28088 | 0.38519 | 0.51085 | 0.64628 | 0.77735 | 0.89303 | 0.98846 | 1.0619 | 1.1159 | 1.155 | 1.183 | 1.203 | 1.2167 | 1.2267 | 1.2373 | 1.2444 | 1.2521 | 1.2603 | 1.2702 | 1.2796 | 1.2944 | 1.3089 | 1.3285 | 1.3531 | 1.3812 | 1.4143 | 1.4538 | 1.499
Time 8.5 h, -40 celcius | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.083695 | 0.08324 | 0.083622 | 0.083847 | 0.084834 | 0.086129 | 0.086714 | 0.087381 | 0.087886 | 0.088558 | 0.089009 | 0.089429 | 0.089952 | 0.090576 | 0.09152 | 0.092833 | 0.094685 | 0.097293 | 0.10083 | 0.10562 | 0.11226 | 0.12153 | 0.13502 | 0.15512 | 0.18502 | 0.22832 | 0.28819 | 0.36571 | 0.45912 | 0.56375 | 0.67242 | 0.77798 | 0.87404 | 0.95702 | 1.0284 | 1.0845 | 1.1268 | 1.1614 | 1.1873 | 1.2052 | 1.2167 | 1.2287 | 1.2359 | 1.2394 | 1.2468 | 1.2512 | 1.2552 | 1.2597
Std. : Water (1:0)_ 5 g_17C | 2.3303 | 2.4387 | 2.4946 | 2.5334 | 2.5734 | 2.595 | 2.6101 | 2.6237 | 2.6363 | 2.6458 | 2.6525 | 2.6605 | 2.6675 | 2.6754 | 2.6841 | 2.6936 | 2.7028 | 2.7121 | 2.7221 | 2.732 | 2.7416 | 2.7519 | 2.7627 | 2.773 | 2.7842 | 2.7949 | 2.8043 | 2.8147 | 2.8246 | 2.8354 | 2.8466 | 2.8598 | 2.8739 | 2.8899 | 2.908 | 2.9285 | 2.9492 | 2.972 | 2.9971 | 3.0212 | 3.0504 | 3.076 | 3.1069 | 3.1381 | 3.1714 | 3.2072 | 3.2459 | 3.2864
Std. : Water (1:3)_5 g_22.3C | 1.4671 | 1.6855 | 1.744 | 1.8824 | 1.9929 | 2.1292 | 2.3085 | 2.4363 | 2.538 | 2.631 | 2.6959 | 2.7426 | 2.7766 | 2.8037 | 2.8219 | 2.8372 | 2.8511 | 2.8646 | 2.876 | 2.8875 | 2.8997 | 2.9118 | 2.9235 | 2.9355 | 2.9478 | 2.9606 | 2.9734 | 2.9866 | 2.9991 | 3.0123 | 3.0249 | 3.039 | 3.0531 | 3.0685 | 3.0846 | 3.1036 | 3.1222 | 3.1442 | 3.1674 | 3.1919 | 3.2199 | 3.2461 | 3.2797 | 3.3073 | 3.3426 | 3.3775 | 3.4158 | 3.4569 | 3.5017 | 3.5489
Fitting-Time 0.23 h (20.3 C) Liquid | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.412 | 2.4119 | 2.4119 | 2.4119 | 2.4119 | 2.4119 | 2.4118 | 2.4117 | 2.4115 | 2.4112 | 2.4108 | 2.4101 | 2.409 | 2.4073 | 2.4046 | 2.4004 | 2.3937 | 2.3832 | 2.367 | 2.342 | 2.3043 | 2.2484 | 2.1697 | 2.0642 | 1.9301 | 1.7781 | 1.6192 | 1.4715 | 1.3445 | 1.2461 | 1.1731 | 1.1243 | 1.0907 | 1.0692 | 1.0535 | 1.0451 | 1.0395 | 1.035 | 1.0329 | 1.0318 | 1.0305 | 1.0302 | 1.0299 | 1.0297
Fitting-Time 5.03 h (-20 C) Frozen | 1.2389 | 1.2389 | 1.2389 | 1.2389 | 1.2388 | 1.2387 | 1.2386 | 1.2384 | 1.2381 | 1.2377 | 1.2369 | 1.2358 | 1.2339 | 1.231 | 1.2264 | 1.2192 | 1.208 | 1.1906 | 1.1642 | 1.1248 | 1.0679 | 0.98926 | 0.88724 | 0.76529 | 0.63274 | 0.50242 | 0.38716 | 0.29408 | 0.22349 | 0.17372 | 0.13979 | 0.11692 | 0.10222 | 0.092682 | 0.086624 | 0.08272 | 0.080267 | 0.078693 | 0.077739 | 0.077121 | 0.076743 | 0.076474 | 0.076332 | 0.076237 | 0.076164 | 0.07613 | 0.076111 | 0.076089 | 0.076084 | 0.076079 | 0.076076
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124112991 | 5.7569630370485525 | 5.6020599913279625 | 5.5228744023138026 | 5.4259741609193757 | 5.3010299956639821 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521985 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639821 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.9004927554293705 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463871 | 3.4001233554370813 | 3.3001605369513531 | 3.2001388538573816 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279803 | 2.7000023971422542 | 2.600024903892987 | 2.5000030680516936 | 2.4000023459279562 | 2.3000082025538129 | 2.2000018654066023 | 2.0999912335446842 | 2 | 1.9000009651534697 | 1.8000018278179688 | 1.7000023971422538 | 1.6000030866277406 | 1.5000030680516931 | 1.4000023459279558 | 1.3000082025538131 | 1.200001865406602 | 1.099991233544684 | 1
Time 0.23 h, 20.3 celcius | 0.021874 | 0.027746 | 0.035469 | 0.045629 | 0.058992 | 0.069416 | 0.087012 | 0.11672 | 0.14702 | 0.18167 | 0.23092 | 0.28348 | 0.3471 | 0.41787 | 0.49639 | 0.57052 | 0.6315 | 0.66609 | 0.66868 | 0.64015 | 0.58783 | 0.52385 | 0.45806 | 0.3962 | 0.3406 | 0.2936 | 0.25428 | 0.22269 | 0.19622 | 0.17597 | 0.15804 | 0.14279 | 0.13021 | 0.12053 | 0.11142 | 0.10593 | 0.099251 | 0.10112 | 0.096409 | 0.099903 | 0.099379 | 0.11049 | 0.12227 | 0.13167 | 0.14829 | 0.16415 | 0.18118 | 0.20571 | 0.22768 | 0.25157 | 0.27443
Time 5.03 h, -20 celcius | 0.0046903 | 0.0060276 | 0.0048957 | 0.0052689 | 0.0080706 | 0.0070995 | 0.0085621 | 0.011385 | 0.014441 | 0.017939 | 0.023068 | 0.028754 | 0.036271 | 0.045614 | 0.058069 | 0.073874 | 0.094576 | 0.12111 | 0.15588 | 0.20027 | 0.2566 | 0.32289 | 0.39693 | 0.47149 | 0.53464 | 0.57592 | 0.58749 | 0.56901 | 0.52696 | 0.47094 | 0.40966 | 0.35018 | 0.29633 | 0.24993 | 0.21217 | 0.18272 | 0.16189 | 0.14493 | 0.13366 | 0.12856 | 0.12752 | 0.12988 | 0.13813 | 0.14441 | 0.16577 | 0.18416 | 0.20379 | 0.22754 | 0.2521 | 0.27626 | 0.29768
Time 8.5 h, -40 celcius | 0.0037257 | 0.0037257 | 0.0022099 | 0.0055149 | 0.0027352 | 0.0046061 | 0.0028107 | 0.0031095 | 0.0040131 | 0.0050574 | 0.0061923 | 0.007861 | 0.0096387 | 0.01196 | 0.014763 | 0.018394 | 0.022849 | 0.028488 | 0.035374 | 0.043956 | 0.054453 | 0.067485 | 0.083212 | 0.10248 | 0.12627 | 0.15518 | 0.19059 | 0.23316 | 0.28257 | 0.33712 | 0.39303 | 0.44437 | 0.48479 | 0.5087 | 0.51347 | 0.50011 | 0.4716 | 0.43276 | 0.38638 | 0.3379 | 0.29199 | 0.24717 | 0.21023 | 0.17522 | 0.14921 | 0.12645 | 0.10764 | 0.094062 | 0.085294 | 0.077832 | 0.076166 | 0.074374
Std. : Water (1:0)_ 5 g_17C | 0.40745 | 0.37758 | 0.33167 | 0.27797 | 0.22443 | 0.19417 | 0.16169 | 0.13849 | 0.12143 | 0.10697 | 0.096405 | 0.087034 | 0.081871 | 0.078656 | 0.076272 | 0.075226 | 0.07448 | 0.073601 | 0.073296 | 0.073701 | 0.074654 | 0.075483 | 0.075353 | 0.07631 | 0.0769 | 0.077719 | 0.079173 | 0.081297 | 0.084848 | 0.090883 | 0.098449 | 0.10715 | 0.11892 | 0.13057 | 0.14131 | 0.15224 | 0.161 | 0.1718 | 0.18315 | 0.19373 | 0.20789 | 0.22266 | 0.23803 | 0.25543 | 0.2743 | 0.29629 | 0.32039 | 0.34659
Std. : Water (1:3)_5 g_22.3C | 0.35022 | 0.37378 | 0.41173 | 0.55328 | 0.59494 | 0.61621 | 0.60526 | 0.57334 | 0.52449 | 0.45124 | 0.3927 | 0.33463 | 0.28521 | 0.24069 | 0.2067 | 0.17971 | 0.15895 | 0.14354 | 0.13098 | 0.1209 | 0.11373 | 0.10865 | 0.10508 | 0.10263 | 0.10026 | 0.099388 | 0.099009 | 0.09943 | 0.10029 | 0.10249 | 0.10653 | 0.11105 | 0.11816 | 0.12566 | 0.13766 | 0.14944 | 0.16023 | 0.17158 | 0.18231 | 0.19465 | 0.20803 | 0.22162 | 0.23412 | 0.2564 | 0.27568 | 0.29801 | 0.32242 | 0.35087 | 0.38221 | 0.41669
Fitting-Time 0.23 h (20.3 C) Liquid | 0.00080148 | 0.001009 | 0.0012703 | 0.0015992 | 0.0020132 | 0.0025345 | 0.0031908 | 0.0040169 | 0.005057 | 0.0063663 | 0.0080146 | 0.010089 | 0.012702 | 0.01599 | 0.020128 | 0.025337 | 0.031892 | 0.040136 | 0.050507 | 0.063538 | 0.07988 | 0.10037 | 0.126 | 0.15786 | 0.1972 | 0.24538 | 0.30315 | 0.37047 | 0.44655 | 0.52567 | 0.59992 | 0.65886 | 0.68897 | 0.68388 | 0.64493 | 0.58011 | 0.50283 | 0.42213 | 0.34988 | 0.28493 | 0.23165 | 0.18163 | 0.14738 | 0.11839 | 0.089082 | 0.071374 | 0.059526 | 0.041708 | 0.035758 | 0.029805 | 0.023848
Fitting-Time 5.03 h (-20 C) Frozen | 0.0048296 | 0.0060799 | 0.0076542 | 0.0096359 | 0.01213 | 0.01527 | 0.019222 | 0.024195 | 0.030452 | 0.038322 | 0.048213 | 0.060635 | 0.076215 | 0.095709 | 0.12001 | 0.15014 | 0.18717 | 0.232 | 0.28517 | 0.34602 | 0.41191 | 0.4775 | 0.53412 | 0.57135 | 0.58091 | 0.56037 | 0.51476 | 0.45386 | 0.38691 | 0.32252 | 0.26465 | 0.2141 | 0.17243 | 0.138 | 0.11029 | 0.087693 | 0.069745 | 0.055173 | 0.044035 | 0.034967 | 0.027983 | 0.021692 | 0.017495 | 0.013998 | 0.010499 | 0.0083994 | 0.0069995 | 0.0048997 | 0.0041998 | 0.0034999 | 0.0027999
Real part Capacitance
Imaginary part Capacitance
(I)
18 GHz
20.3 °C
20.3 °C
18 kHz
-20 °C
(III)
2.6 kHz
-20 °C
-40 °C
-40 °C
537 Hz
(IV)
(II)
0.0056 M KCl
0.0056 M KCl
Notes: D:\PhD\Assignment\Conference ppt\Plot used in smi 2019 Conference_v1

## Slide 12: Further Reading
Introduction to TVIS theory
Description of the measurement principles
Dielectric loss and relaxations mechanisms (liquid and frozen states)

## Slide 13: Dielectric loss or dielectric permittivity analysis?
Through Vial Impedance Spectroscopy (TVIS)
Dielectric loss or dielectric permittivity analysis?

## Slide 14: Applications for the dielectric loss spectrum
Chart: Log Frequency
| 1 | 1.0999912335446844 | 1.2000018654066018 | 1.3000082025538131 | 1.400002345927956 | 1.5000030680516931 | 1.6000030866277404 | 1.700002397142254 | 1.8000018278179688 | 1.9000009651534695 | 2 | 2.0999912335446846 | 2.2000018654066018 | 2.3000082025538129 | 2.4000023459279558 | 2.5000030680516931 | 2.6000249038929866 | 2.7000023971422538 | 2.8000362418279794 | 2.9000611027354806 | 3 | 3.1000257301078626 | 3.2001388538573807 | 3.3001605369513523 | 3.4001233554370809 | 3.5003188237463867 | 3.6003302305807652 | 3.7000543856282389 | 3.800655228145212 | 3.90049275542937 | 4 | 4.1009905230699646 | 4.200659016251719 | 4.3010299956639813 | 4.4006588497649739 | 4.5016890119157003 | 4.6020599913279625 | 4.7044336600521994 | 4.8027190075796815 | 4.9030899869919438 | 5 | 5.1106906972737542 | 5.204119982655925 | 5.3010299956639813 | 5.4259741609193757 | 5.5228744023138034 | 5.6020599913279625 | 5.7569630370485534 | 5.8239109124113 | 5.9030899869919438 | 6
Time 13.1 h | 0.10929 | 0.1004 | 0.092502 | 0.087998 | 0.086049 | 0.088115 | 0.0938 | 0.10112 | 0.11429 | 0.13106 | 0.15364 | 0.18166 | 0.21583 | 0.25512 | 0.29985 | 0.34774 | 0.3951 | 0.43795 | 0.46981 | 0.48496 | 0.48002 | 0.45469 | 0.41262 | 0.36037 | 0.30465 | 0.25101 | 0.20331 | 0.16311 | 0.13007 | 0.10386 | 0.083113 | 0.06641 | 0.053298 | 0.042748 | 0.034381 | 0.027604 | 0.022219 | 0.017839 | 0.014449 | 0.011634 | 0.0093969 | 0.0073773 | 0.0059706 | 0.0046834 | 0.003604 | 0.003119 | 0.0043457 | 0.0022557 | 0.0032184 | 0.0030577 | 0.0023234
Time 16.1 h | 0.11007 | 0.10093 | 0.090991 | 0.082047 | 0.079543 | 0.078788 | 0.081723 | 0.087173 | 0.096992 | 0.11005 | 0.12829 | 0.15003 | 0.17689 | 0.20803 | 0.24102 | 0.27482 | 0.30518 | 0.32877 | 0.34107 | 0.33971 | 0.32434 | 0.2969 | 0.26113 | 0.22177 | 0.1829 | 0.1475 | 0.11733 | 0.092752 | 0.073068 | 0.057802 | 0.045912 | 0.036469 | 0.029123 | 0.02329 | 0.018706 | 0.015032 | 0.012128 | 0.0097734 | 0.0079623 | 0.0064562 | 0.0052363 | 0.0041244 | 0.0033647 | 0.0027806 | 0.0022635 | 0.0029162 | 0.0030408 | 0.0014411 | 0.0021872 | 0.0031647 | 0.0015537
Time 19.1 h | 0.096631 | 0.084665 | 0.074113 | 0.066591 | 0.062691 | 0.058864 | 0.058511 | 0.060476 | 0.065071 | 0.072096 | 0.081916 | 0.094578 | 0.10999 | 0.12716 | 0.1449 | 0.16219 | 0.17646 | 0.18582 | 0.1881 | 0.18242 | 0.16923 | 0.15032 | 0.12827 | 0.10587 | 0.085142 | 0.067241 | 0.052563 | 0.040958 | 0.031878 | 0.024951 | 0.019669 | 0.015531 | 0.012366 | 0.0098704 | 0.0079292 | 0.0063816 | 0.0051741 | 0.0042062 | 0.0034663 | 0.0028605 | 0.0023538 | 0.0018919 | 0.0016185 | 0.0014417 | 0.0019173 | 0.0019251 | 0.0021233 | 0.0015948 | 0.0015646 | 0.0023139 | 0.00072884
 | 0.11394 | 0.093481 | 0.077657 | 0.065696 | 0.057013 | 0.051186 | 0.047934 | 0.047103 | 0.048651 | 0.05264 | 0.059221 | 0.068608 | 0.08103 | 0.096621 | 0.11525 | 0.13624 | 0.15804 | 0.17801 | 0.19272 | 0.1989 | 0.19496 | 0.18175 | 0.16215 | 0.13967 | 0.1172 | 0.096485 | 0.078464 | 0.063316 | 0.050716 | 0.040549 | 0.032372 | 0.025721 | 0.020479 | 0.01627 | 0.012943 | 0.010261 | 0.008146 | 0.0064366 | 0.0051337 | 0.0040747 | 0.00326 | 0.0025268 | 0.0020378 | 0.0016303
 | 0.12912 | 0.10754 | 0.091277 | 0.079538 | 0.071764 | 0.067588 | 0.066822 | 0.06944 | 0.075575 | 0.085504 | 0.099635 | 0.11846 | 0.1425 | 0.17204 | 0.20691 | 0.24592 | 0.28627 | 0.32317 | 0.35038 | 0.36195 | 0.35498 | 0.33106 | 0.29542 | 0.25449 | 0.21357 | 0.17583 | 0.143 | 0.11539 | 0.092429 | 0.0739 | 0.058998 | 0.046874 | 0.037321 | 0.02965 | 0.023587 | 0.018699 | 0.014845 | 0.01173 | 0.0093551 | 0.0074254 | 0.0059407 | 0.0046044 | 0.0037133 | 0.0029707
 | 0.1177 | 0.10593 | 0.097045 | 0.090946 | 0.087643 | 0.087239 | 0.089948 | 0.096102 | 0.10617 | 0.12074 | 0.14056 | 0.16644 | 0.1992 | 0.2394 | 0.28694 | 0.3404 | 0.39622 | 0.44803 | 0.48732 | 0.50564 | 0.49835 | 0.46702 | 0.41851 | 0.36179 | 0.30443 | 0.25118 | 0.20462 | 0.16535 | 0.13261 | 0.10615 | 0.084841 | 0.067487 | 0.053798 | 0.042795 | 0.034092 | 0.027068 | 0.021523 | 0.017036 | 0.013612 | 0.010826 | 0.00868 | 0.0067451 | 0.0054527 | 0.0043739
Drying time
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.003361 | 0.0040465 | 0.0025946 | 0.0026214 | 0.0044553 | 0.0026323 | 0.0029087 | 0.0037366 | 0.0047149 | 0.0057564 | 0.0073092 | 0.0089323 | 0.011077 | 0.013703 | 0.017072 | 0.021236 | 0.026498 | 0.03301 | 0.041177 | 0.051293 | 0.064055 | 0.079683 | 0.099225 | 0.12384 | 0.15437 | 0.19231 | 0.23844 | 0.29224 | 0.35137 | 0.41125 | 0.46468 | 0.50437 | 0.52462 | 0.52356 | 0.50251 | 0.46708 | 0.42347 | 0.37253 | 0.32299 | 0.27537 | 0.23266 | 0.1978 | 0.16633 | 0.14424 | 0.12252 | 0.10889 | 0.097697 | 0.093226 | 0.09166 | 0.092155 | 0.093069
Time 4.5 h | 0.0034249 | 0.0041377 | 0.0027283 | 0.0027814 | 0.0046887 | 0.0029259 | 0.0032733 | 0.0042363 | 0.0053358 | 0.0065296 | 0.0082746 | 0.010141 | 0.012601 | 0.015587 | 0.01947 | 0.024233 | 0.030264 | 0.037737 | 0.047171 | 0.058901 | 0.073903 | 0.092482 | 0.11581 | 0.14538 | 0.18208 | 0.2274 | 0.28124 | 0.3416 | 0.40418 | 0.46243 | 0.50811 | 0.53414 | 0.53731 | 0.51809 | 0.48197 | 0.43577 | 0.38293 | 0.3311 | 0.28177 | 0.23748 | 0.20077 | 0.16802 | 0.14416 | 0.11848 | 0.11171 | 0.10187 | 0.09506 | 0.092791 | 0.094429 | 0.097904 | 0.10371
Time 4.6 h | 0.0036091 | 0.0043103 | 0.0029276 | 0.0030069 | 0.0049863 | 0.0032919 | 0.0037411 | 0.0048548 | 0.006126 | 0.0075084 | 0.0095269 | 0.011725 | 0.014579 | 0.018062 | 0.02259 | 0.028182 | 0.035264 | 0.044078 | 0.055262 | 0.069282 | 0.08729 | 0.10998 | 0.13862 | 0.17503 | 0.21981 | 0.27381 | 0.33549 | 0.40061 | 0.46249 | 0.51282 | 0.5432 | 0.54939 | 0.53148 | 0.49439 | 0.44517 | 0.39079 | 0.33577 | 0.28667 | 0.23998 | 0.20372 | 0.17181 | 0.14592 | 0.12721 | 0.11818 | 0.10489 | 0.098824 | 0.097908 | 0.099283 | 0.10394 | 0.11202 | 0.12205
Time 4.7 h | 0.0040254 | 0.0023197 | 0.0060162 | 0.0032852 | 0.0053707 | 0.003773 | 0.0043423 | 0.005672 | 0.0071637 | 0.0088006 | 0.011181 | 0.013771 | 0.017191 | 0.021329 | 0.026752 | 0.033478 | 0.04203 | 0.052722 | 0.066401 | 0.083738 | 0.10613 | 0.13436 | 0.17037 | 0.21573 | 0.27035 | 0.33347 | 0.40107 | 0.46603 | 0.51937 | 0.55253 | 0.55996 | 0.54171 | 0.50311 | 0.45097 | 0.39455 | 0.33876 | 0.28887 | 0.24104 | 0.20387 | 0.17339 | 0.14886 | 0.13225 | 0.11753 | 0.11 | 0.10482 | 0.10438 | 0.1074 | 0.1138 | 0.12408 | 0.13639 | 0.14993
Time 4.8 h | 0.0039527 | 0.0047899 | 0.0034912 | 0.0036734 | 0.0059298 | 0.004463 | 0.0052209 | 0.0068624 | 0.0086692 | 0.010687 | 0.013626 | 0.016821 | 0.021041 | 0.026213 | 0.03302 | 0.041432 | 0.052261 | 0.065911 | 0.083539 | 0.10606 | 0.13545 | 0.17252 | 0.21928 | 0.27665 | 0.34237 | 0.41271 | 0.47977 | 0.53353 | 0.56486 | 0.56874 | 0.5461 | 0.50317 | 0.44813 | 0.38872 | 0.33224 | 0.28045 | 0.23905 | 0.20172 | 0.17254 | 0.14998 | 0.13098 | 0.12232 | 0.11386 | 0.11155 | 0.11412 | 0.11812 | 0.12658 | 0.13848 | 0.15361 | 0.17099 | 0.18866
Time 4.9 h | 0.0043188 | 0.0052277 | 0.0039868 | 0.0042288 | 0.0066529 | 0.0053479 | 0.0063485 | 0.0083796 | 0.010604 | 0.013107 | 0.016765 | 0.020805 | 0.02609 | 0.032614 | 0.041224 | 0.052023 | 0.065957 | 0.083695 | 0.10683 | 0.13651 | 0.17523 | 0.22353 | 0.28261 | 0.35128 | 0.42376 | 0.49222 | 0.54583 | 0.5749 | 0.57491 | 0.54814 | 0.50162 | 0.4443 | 0.38384 | 0.32634 | 0.27564 | 0.23272 | 0.19908 | 0.1717 | 0.1495 | 0.13558 | 0.12502 | 0.11938 | 0.11838 | 0.12437 | 0.13062 | 0.13992 | 0.15492 | 0.1723 | 0.19073 | 0.21306 | 0.23377
Time 5 h | 0.0049133 | 0.0058042 | 0.0046338 | 0.0049712 | 0.0076664 | 0.0066031 | 0.0079479 | 0.010533 | 0.013358 | 0.01657 | 0.021298 | 0.026516 | 0.033397 | 0.041936 | 0.053268 | 0.067599 | 0.086277 | 0.11024 | 0.14161 | 0.18183 | 0.2334 | 0.29549 | 0.36662 | 0.44167 | 0.50969 | 0.5605 | 0.58424 | 0.57776 | 0.54458 | 0.49349 | 0.43374 | 0.3731 | 0.31662 | 0.26729 | 0.22557 | 0.19336 | 0.16657 | 0.14984 | 0.13717 | 0.1291 | 0.12498 | 0.12759 | 0.13258 | 0.1444 | 0.15601 | 0.17194 | 0.19151 | 0.21398 | 0.23437 | 0.26066 | 0.28203
Time 5.1 h | 0.0051615 | 0.0065394 | 0.0054847 | 0.0059393 | 0.0090001 | 0.0082333 | 0.010001 | 0.013324 | 0.016914 | 0.021051 | 0.027147 | 0.03396 | 0.042946 | 0.054219 | 0.069319 | 0.088593 | 0.11394 | 0.14657 | 0.18909 | 0.24251 | 0.30812 | 0.38167 | 0.45767 | 0.52576 | 0.57299 | 0.59132 | 0.57824 | 0.5397 | 0.48481 | 0.42344 | 0.36292 | 0.3071 | 0.25963 | 0.22024 | 0.18846 | 0.16573 | 0.14947 | 0.13846 | 0.13032 | 0.13113 | 0.1334 | 0.1422 | 0.15348 | 0.16641 | 0.18887 | 0.2096 | 0.23238 | 0.25699 | 0.28353 | 0.3063 | 0.32649
-38 oC
-18 oC
temperature calibration for predicting temperature of the product in primary drying
drying rate
from
Chart: Temperature/°C
| -38.4 | -37.299999999999997 | -36.5 | -35.700000000000003 | -34.799999999999997 | -33.799999999999997 | -32.799999999999997 | -32.1 | -31 | -30.2 | -28.9 | -27.9 | -27 | -26 | -25.1 | -24 | -23 | -22 | -21.1 | -20 | -19.100000000000001 | -18.2
Log(Frequency) | 2.7551 | 2.7816 | 2.8077 | 2.8379 | 2.8682 | 2.9001 | 2.9302 | 2.9619 | 2.9942 | 3.0287 | 3.0652 | 3.1011 | 3.139 | 3.177 | 3.2151 | 3.2535 | 3.2935 | 3.3335 | 3.3735 | 3.4135 | 3.4535 | 3.4959
C''PEAK | 0.51695 | 0.51973 | 0.52247 | 0.52612 | 0.52968 | 0.53324 | 0.53671 | 0.54012 | 0.54345 | 0.54669 | 0.54988 | 0.55289 | 0.5558 | 0.55853 | 0.56117 | 0.5636 | 0.56593 | 0.56819 | 0.57022 | 0.57217 | 0.57397 | 0.5757
Temperature compensation for prior to determination of drying rate
These concepts were used in our recent paper :
Smith, G., Jeeraruangrattana, Y., Ermolina, I. (2018). The application of dual-electrode through vial impedance spectroscopy for the determination of ice interface temperatures, primary drying rate and vial heat transfer coefficient in lyophilization process development. European Journal of Pharmaceutics and Biopharmaceutics.
Notes: Eg temperature cycling below to calibrate Fpeak for temperature (in order to predict product temperatures in primary drying) and the compensate cpeak for temperature so that it can used to monitor ice content and hence drying rate
Eg 2 left graph: monitoring phase behaviour (ice nucleation temperatures and solidufication end points)

## Slide 15: Applications for dielectric permittivity spectrum
is highly sensitive to low ice volumes
To date we have been using that for the determination end point of primary drying. # See Conference Poster 11 (Bhaskar et al)
More recently we have started using the dielectric permittivity spectrum for
  Ice nucleation temperatures
  Ice crystallization end-point
  Glass transition temperature………………” The focus for the rest of the presentation”
https://doi.org/10.21253/DMU.9771695
Notes: D:\PhD\Experiment\2016_06_29_Full Shelf Water Study\281016 (Different height from base)\3 mm from base

## Slide 16: Dielectric Permittivity Spectrum: What frequency?
Through Vial Impedance Spectroscopy (TVIS)
Dielectric Permittivity Spectrum: What frequency?

## Slide 17: Applications for dielectric permittivity spectrum
Temperature sensitivity of the real part capacitance (dielectric storage or dielectric permittivity) of the TVIS vial (containing ice) depends on the measurement frequency
The low frequency capacitance is strongly temperature dependent
The high frequency capacitance is weakly temperature dependent
Chart: Temperature/°C
| -38.4 | -37.299999999999997 | -36.5 | -35.700000000000003 | -34.799999999999997 | -33.799999999999997 | -32.799999999999997 | -32.1 | -31 | -30.2 | -28.9 | -27.9 | -27 | -26 | -25.1 | -24 | -23 | -22 | -21.1 | -20 | -19.100000000000001 | -18.2
Tslice (10 Hz) | 1 | 1.002 | 1.0053 | 1.0078 | 1.0124 | 1.017 | 1.0218 | 1.028 | 1.0342 | 1.0419 | 1.0517 | 1.0616 | 1.0743 | 1.0873 | 1.1018 | 1.1166 | 1.1355 | 1.1535 | 1.173 | 1.1944 | 1.2161 | 1.2383
Tslice (100 Hz) | 1 | 1.0081 | 1.0133 | 1.0204 | 1.0254 | 1.0305 | 1.0348 | 1.0392 | 1.0435 | 1.0473 | 1.0504 | 1.0541 | 1.0592 | 1.0616 | 1.065 | 1.0688 | 1.0733 | 1.0764 | 1.0803 | 1.0843 | 1.0884 | 1.0927
Tslice (0.2MHz) | 1 | 0.99993 | 0.99767 | 1.0006 | 1.0015 | 1.0013 | 1.0025 | 1.0029 | 1.0036 | 1.004 | 1.0039 | 1.0045 | 1.0048 | 1.0061 | 1.0056 | 1.0058 | 1.0069 | 1.007 | 1.0082 | 1.0086 | 1.0087 | 1.0086
Chart: Log Frequency
| 6 | 5.9030899869919438 | 5.8239109124113 | 5.7569630370485534 | 5.6020599913279625 | 5.5228744023138034 | 5.4259741609193757 | 5.3010299956639813 | 5.204119982655925 | 5.1106906972737542 | 5 | 4.9030899869919438 | 4.8027190075796815 | 4.7044336600521994 | 4.6020599913279625 | 4.5016890119157003 | 4.4006588497649739 | 4.3010299956639813 | 4.200659016251719 | 4.1009905230699646 | 4 | 3.90049275542937 | 3.800655228145212 | 3.7000543856282389 | 3.6003302305807652 | 3.5003188237463867 | 3.4001233554370809 | 3.3001605369513523 | 3.2001388538573807 | 3.1000257301078626 | 3 | 2.9000611027354806 | 2.8000362418279794 | 2.7000023971422538 | 2.6000249038929866 | 2.5000030680516931 | 2.4000023459279558 | 2.3000082025538129 | 2.2000018654066018 | 2.0999912335446846 | 2 | 1.9000009651534695 | 1.8000018278179688 | 1.700002397142254 | 1.6000030866277404 | 1.5000030680516931 | 1.400002345927956 | 1.3000082025538131 | 1.2000018654066018 | 1.0999912335446844 | 1
Time 4.4 h | 0.084931 | 0.086522 | 0.084416 | 0.083584 | 0.083174 | 0.083511 | 0.083751 | 0.08481 | 0.086097 | 0.087016 | 0.08765 | 0.088181 | 0.088496 | 0.088762 | 0.089099 | 0.089472 | 0.089925 | 0.090513 | 0.091358 | 0.092525 | 0.094183 | 0.096518 | 0.099835 | 0.10479 | 0.11237 | 0.12448 | 0.14398 | 0.17467 | 0.22057 | 0.28526 | 0.36955 | 0.47045 | 0.58177 | 0.69523 | 0.80256 | 0.89741 | 0.97772 | 1.0421 | 1.0945 | 1.1331 | 1.1623 | 1.1841 | 1.1989 | 1.2068 | 1.2187 | 1.2253 | 1.2327 | 1.237 | 1.2399 | 1.2486 | 1.255
Time 4.5 h | 0.085366 | 0.086803 | 0.084615 | 0.083791 | 0.08325 | 0.083655 | 0.08388 | 0.084857 | 0.086134 | 0.087029 | 0.087739 | 0.088261 | 0.088595 | 0.088921 | 0.089249 | 0.089597 | 0.090054 | 0.090652 | 0.091444 | 0.092572 | 0.094147 | 0.096469 | 0.1002 | 0.10625 | 0.11628 | 0.13294 | 0.16019 | 0.20224 | 0.26295 | 0.34453 | 0.4451 | 0.55868 | 0.67675 | 0.78985 | 0.89073 | 0.97528 | 1.0447 | 1.0963 | 1.1358 | 1.1645 | 1.186 | 1.201 | 1.2113 | 1.2236 | 1.2271 | 1.2334 | 1.2393 | 1.2455 | 1.2503 | 1.2572 | 1.2649
Time 4.6 h | 0.085421 | 0.086928 | 0.084724 | 0.083845 | 0.083347 | 0.08379 | 0.084043 | 0.085026 | 0.086304 | 0.087178 | 0.087815 | 0.088397 | 0.088756 | 0.089066 | 0.089366 | 0.089722 | 0.09011 | 0.090631 | 0.091373 | 0.092445 | 0.094129 | 0.096729 | 0.10147 | 0.10987 | 0.1245 | 0.14913 | 0.18836 | 0.24664 | 0.32654 | 0.42723 | 0.54333 | 0.66518 | 0.78294 | 0.88848 | 0.97692 | 1.0466 | 1.0994 | 1.1385 | 1.1666 | 1.1884 | 1.2027 | 1.2152 | 1.2223 | 1.2362 | 1.2355 | 1.2434 | 1.246 | 1.2536 | 1.263 | 1.2718 | 1.2824
Time 4.7 h | 0.085219 | 0.053673 | 0.07561 | 0.08397 | 0.083322 | 0.083877 | 0.08413 | 0.085153 | 0.086388 | 0.087288 | 0.087947 | 0.088541 | 0.088887 | 0.089203 | 0.089467 | 0.089732 | 0.090058 | 0.090474 | 0.091108 | 0.092199 | 0.094218 | 0.097969 | 0.10503 | 0.11817 | 0.14112 | 0.17863 | 0.23583 | 0.31561 | 0.41746 | 0.5363 | 0.66212 | 0.78332 | 0.89159 | 0.98177 | 1.0528 | 1.1059 | 1.144 | 1.1727 | 1.1928 | 1.2063 | 1.2173 | 1.2263 | 1.2337 | 1.2352 | 1.2458 | 1.2532 | 1.2583 | 1.2687 | 1.2782 | 1.2933 | 1.3076
Time 4.8 h | 0.085767 | 0.086823 | 0.084565 | 0.083702 | 0.083251 | 0.083937 | 0.084213 | 0.085213 | 0.08646 | 0.08732 | 0.087939 | 0.088605 | 0.088971 | 0.089203 | 0.089353 | 0.089534 | 0.089704 | 0.089999 | 0.090613 | 0.092049 | 0.095196 | 0.10173 | 0.11443 | 0.13773 | 0.1764 | 0.23592 | 0.3197 | 0.42654 | 0.54993 | 0.67941 | 0.80269 | 0.91088 | 0.99931 | 1.0678 | 1.1184 | 1.1539 | 1.1788 | 1.1983 | 1.2124 | 1.2227 | 1.2311 | 1.2384 | 1.2447 | 1.2451 | 1.2597 | 1.2688 | 1.2791 | 1.2911 | 1.3078 | 1.3258 | 1.3483
Time 4.9 h | 0.085491 | 0.087251 | 0.08505 | 0.084128 | 0.083174 | 0.083948 | 0.084267 | 0.085306 | 0.08667 | 0.087534 | 0.088121 | 0.088619 | 0.088926 | 0.089085 | 0.089128 | 0.089112 | 0.089161 | 0.08943 | 0.090419 | 0.093078 | 0.099266 | 0.11187 | 0.13539 | 0.17568 | 0.23771 | 0.32517 | 0.43674 | 0.56448 | 0.69641 | 0.82031 | 0.92766 | 1.014 | 1.08 | 1.1283 | 1.1629 | 1.1868 | 1.205 | 1.2169 | 1.2264 | 1.2348 | 1.2422 | 1.2502 | 1.2567 | 1.2693 | 1.277 | 1.2885 | 1.3045 | 1.3218 | 1.3429 | 1.3705 | 1.4013
Time 5 h | 0.085899 | 0.087566 | 0.085217 | 0.084216 | 0.083061 | 0.084092 | 0.084384 | 0.085508 | 0.086793 | 0.087615 | 0.088068 | 0.088506 | 0.088675 | 0.088682 | 0.088573 | 0.088418 | 0.088473 | 0.089248 | 0.091786 | 0.098085 | 0.11166 | 0.13702 | 0.18042 | 0.24808 | 0.34172 | 0.45928 | 0.59178 | 0.72515 | 0.84771 | 0.95135 | 1.0334 | 1.0952 | 1.14 | 1.172 | 1.1948 | 1.2102 | 1.2225 | 1.2305 | 1.2409 | 1.247 | 1.2556 | 1.2643 | 1.2739 | 1.2804 | 1.3003 | 1.3163 | 1.3401 | 1.3649 | 1.3972 | 1.4316 | 1.4722
Time 5.1 h | 0.086471 | 0.087693 | 0.085343 | 0.084296 | 0.082843 | 0.084117 | 0.084511 | 0.08554 | 0.086831 | 0.08761 | 0.088012 | 0.088277 | 0.088283 | 0.088083 | 0.087726 | 0.087608 | 0.088197 | 0.090659 | 0.097298 | 0.11173 | 0.13971 | 0.18675 | 0.25888 | 0.35922 | 0.48223 | 0.61799 | 0.7525 | 0.87288 | 0.97254 | 1.0503 | 1.1079 | 1.1501 | 1.1795 | 1.2007 | 1.2162 | 1.2273 | 1.2375 | 1.2456 | 1.2527 | 1.2602 | 1.27 | 1.2804 | 1.294 | 1.3078 | 1.3301 | 1.3551 | 1.3841 | 1.417 | 1.4564 | 1.5034 | 1.5541
Re-heating of ice
−18 °C
−38 °C
0.2 MHz
10 Hz
10 Hz
100 Hz
0.2 MHz
100 Hz
Notes: D:\PhD\Experiment\2016_06_29_Full Shelf Water Study\281016 (Different height from base)\3 mm from base

## Slide 18: Ice solidification rate
TVIS Applications
Ice solidification rate
Significance: 𝑂𝑝𝑡𝑖𝑚𝑖𝑧𝑎𝑡𝑖𝑜𝑛 𝑜𝑓 𝑖𝑐𝑒 𝑐𝑟𝑦𝑠𝑡𝑎𝑙 𝑠𝑡𝑟𝑢𝑐𝑡𝑢𝑟𝑒𝑠 𝑤𝑖𝑡ℎ l𝑎𝑟𝑔𝑒𝑟 𝑖𝑛𝑡𝑒𝑟𝑐𝑜𝑛𝑛𝑒𝑐𝑡𝑒𝑑 𝑐𝑟𝑦𝑠𝑡𝑎𝑙𝑠 𝑖𝑛𝑐𝑟𝑒𝑎𝑠𝑒𝑠 𝑡ℎ𝑒 𝑝𝑜𝑟𝑜𝑠𝑖𝑡𝑦 𝑜𝑓 𝑡ℎ𝑒 𝑑𝑟𝑦 𝑙𝑎𝑦𝑒𝑟, 𝑤ℎ𝑖𝑐ℎ 𝑖𝑠 𝑡ℎ𝑒 𝑙𝑎𝑦𝑒𝑟 𝑡ℎ𝑎𝑡 𝑖𝑠 𝑟𝑒𝑠𝑡𝑟𝑖𝑐𝑡𝑖𝑛𝑔 𝑡ℎ𝑒 𝑑𝑖𝑓𝑓𝑢𝑠𝑖𝑜𝑛 𝑜𝑓 𝑤𝑎𝑡𝑒𝑟 𝑣𝑎𝑝𝑜𝑢𝑟 fro𝑚 𝑡ℎ𝑒 𝑖𝑐𝑒 𝑖𝑛𝑡𝑒𝑟𝑓𝑎𝑐𝑒

## Slide 19: Nucleation onset
Chart: C′(10 Hz)/pF
Chart: Time/h
Ice nucleation
The capacitance of ice at frequencies below the relaxation frequency of ice (e.g. 10 Hz) is strongly dependent on temperature
Any changes in 10 Hz, either with time or temperature, can be associated with the onset of ice nucleation (which is an exothermic event)

## Slide 20: Ice formation end point
Chart: C′(10 Hz)/pF
Chart: Time/h
Ice solidification end-point
The capacitance of ice has almost no temperature dependence at frequencies above the relaxation frequency of ice (~1 kHz) such as (0.2 MHz).
Any changes in (0.2 MHz) either with time or temperature, can be associated with the completion of ice formation on freezing

## Slide 21: Ice crystallization period
Chart: C′(10 Hz)/pF
Chart: Time/h
Ice nucleation
Solidification end point
The difference between these two times if the ice solidification time
Knowing the height of the product in the vial one can then estimate an average solidification rate
~18 mins
Ice solidification time

## Slide 22: Examples (Edge Vials)
5%Sucrose in 0.55% NaCl
Chart: C′(10 Hz)/pF
Chart: Time/h
1.83 h (-11.9 °C)
36 min
Chart: C′(10 Hz)/pF
Chart: Time/h
5%Sucrose
1.97 h (-10.5 °C)
2.27 h
18 min
Chart: C′ (10 Hz)/pF
Chart: Time/h
1.1 h (-3°C)
1.4 h
18 min
IgG in Mannitol and Sucrose based formulation
Chart: C′ (10 Hz)/pF
Chart: Time/h
IgG in Sucrose based formulation
1.00 h (0.6 °C)
1.29h
17 min
2.44 h

## Slide 23: Determination of Ice Nucleation Temperature ()
TVIS Applications
Determination of Ice Nucleation Temperature ()

## Slide 24: Ice Nucleation Temperature
Chart: Time / h
TVIS record every 120 seconds
(2nd spectrum start after 1st spectrum 10 seconds)
Thermocouple
TC at the middle of solution bounded within electrode region
(defined as Temperature node)
TVIS vial
TC vial
Thermocouple position
Freezing step
Shelf Temperature
Freezing from 20 °C to -45 °C with 0.5 °C/min
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 25: Ice Nucleation Temperature
In case the TVIS vial nucleates before TC vial, the nucleation temperature in the TVIS vial can be inferred directly from TC temperatures in the nearest neighbor vials
Thermocouple
Shelf Temperature
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 26: Ice Nucleation Temperature
Chart: Time / h
In case the TVIS vial nucleates before TC vial, the nucleation temperature in the TVIS vial can be inferred directly from TC temperatures in the nearest neighbor vials
However, if TVIS vial nucleates later than TC vial, the nucleation temperature can be predicted by fitting a curve to the plot of the average temperature from thermocouple vials against TVIS parameter (i.e. )
Thermocouple
Shelf Temperature
Chart: Temperature/°C
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 27: Ice Nucleation Temperature
Chart: Time / h
Chart: C′(10 Hz) / pF
| 3.1565509999999999 | 3.1534010000000001 | 3.1511719999999999 | 3.141886 | 3.129775 | 3.1126879999999999 | 3.0977380000000001 | 3.0769570000000002 | 3.0589249999999999 | 3.0362629999999999 | 3.018599 | 2.9947620000000001 | 2.9750100000000002 | 2.9516149999999999 | 2.9309449999999999 | 2.907899 | 2.8865430000000001 | 2.864188 | 2.8454359999999999 | 2.8234750000000002 | 2.8035779999999999 | 2.7832219999999999 | 2.7638419999999999 | 2.7446600000000001 | 2.7265640000000002 | 2.7092320000000001 | 2.6931080000000001 | 2.674191 | 2.6512959999999999 | 2.6325789999999998 | 2.616924 | 2.6025390000000002 | 2.5895869999999999 | 2.5756290000000002 | 2.5644809999999998 | 2.5525229999999999 | 2.5422660000000001 | 2.532915 | 2.5245259999999998 | 2.519263 | 2.5138340000000001 | 2.5079039999999999 | 2.5022180000000001
Tn from C(10 Hz)-S1 | 19.9 | 19.8 | 19.7 | 19.3 | 19 | 18.7 | 18.2 | 17.8 | 17.2 | 16.6 | 15.9 | 15.3 | 14.7 | 13.9 | 13.2 | 12.5 | 11.8 | 10.9 | 10.3 | 9.4 | 8.6 | 7.9 | 7 | 6.2 | 5.4 | 4.7 | 3.9 | 3.2 | 2.4 | 1.8 | 1.1 | 0.3 | -0.7 | -1.8 | -2.7 | -3.7 | -4.6 | -5.6 | -6.4 | -7.3 | -8 | -8.7 | -9.5
Tn from C(10 Hz)-S1 Trendline | 0.3 | -0.7 | -1.8 | -2.7 | -3.7 | -4.6 | -5.6 | -6.4 | -7.3 | -8 | -8.7 | -9.5
In case the TVIS vial nucleates before TC vial, the nucleation temperature in the TVIS vial can be inferred directly from TC temperatures in the nearest neighbor vials
However, if TVIS vial nucleates later than TC vial, the nucleation temperature can be predicted by fitting a curve to the plot of the average temperature from thermocouple vials against TVIS parameter (i.e. )
The ice nucleation temperature of sample (5 %w/v sucrose) was found to be -10.5 C in the case of this particular TVIS vial (other vials will differ owing to the stochastic nature of ice formation.
Thermocouple
Shelf Temperature
Temperature prediction
Chart: Temperature/°C
−10.5 °C
2.49 pF
Blue region
shows
Pre-nucleation
data
Notes: E:\PhD\Experiment\2018_11_08_Sucrose and NaCl solutions study (Tg' and collapse)

## Slide 28: Conclusions : Ice formation stage
Ice nucleation onset ()
  determined at low frequency (e.g. 100 Hz)
Ice solidification end point ()
  determined at high frequency (e.g. 100 kHz)
Ice solidification time () is the difference between and
Average ice growth rate determined by
Nucleation temperature ()
  determined from extrapolation of pre-nucleation data

## Slide 29: New Book
Chapter 5 Through Vial Impedance Spectroscopy (TVIS) A New Method for Determining the Ice Nucleation Temperature and the Solidification End point

## Slide 30
Jeeraruangrattana et al
# Poster 10
ice formation time and nucleation temperature for Sucrose + NaCl solutions
https://doi.org/10.21253/DMU.9771512

## Slide 31: Determination of in-vial Eutectic melting () or Glass Transition temperature ()
TVIS Applications
Determination of in-vial
Eutectic melting () or
Glass Transition temperature ()
Significance: or underpins/defines the collapse temperature which in turn defines the highest permissible product temperature during primary drying and therefore impacts the maximal achievable drying rate

## Slide 32: PAT for critical temperature determination Tm, Te, Tg
Collapse temperature (defined by formulation and related to Teu, Tg)
  Maximal permissible temperature avoiding structural changes to the product
Lyotherm – integrated electrical Impedance (Zsinφ) and DTA
designed to measure glass transition (Tg’), eutectic (Teu) and melting (Tm) temperatures relevant to freeze-drying formulations
Ward & Matejtschuk , 2010 in Freeze Drying/ Lyophilization of Pharmaceutical & Biological Products 3rd ed. Rey,L & May JC eds, Informa Press, New York

## Slide 33: Glass Transition Temperature
Chart: Time /h
TVIS spectrum recorded every 2 min
Thermocouple
TC at the middle of solution bounded within electrode region
(defined as Temperature node)
TVIS vial
TC vial
Thermocouple position
Re-heating of pre-frozen solution
Ramp from -50 °C to -10 °C at 0.2 /min
Shelf Temperature
Notes: E:\PhD\Experiment\2017_11_01_Glass Transition Study of Sugar\171117-Sucrose

## Slide 34: IgG formulations : melt back vs glass transition
Chart: Temperature/°C
Tslice ( Hz) | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Chart: Temperature/°C
Tslice ( Hz) | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Chart: Temperature/°C
Tslice ( Hz) | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Chart: Temperature/°C
Tslice ( Hz) | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Mannitol+Sucrose based formulation at Edge
Mannitol+Sucrose based formulation
at Core
Sucrose based formulation
at Edge
Sucrose based formulation
at Core
Mannitol
melting
= −34 °C

## Slide 35: Conclusions concerning real part capacitance spectrum
Low frequency dielectric properties of ice
  Pronounced temperature dependency
  Determination of the onset of ice formation
  (and time point when excess thermal energy has dissipated from the system – use in defining start of annealing phase)
High frequency dielectric properties of ice
  Negligible temperature dependency
  Determination of end point of ice crystallization
  Mono-tonic changes with product temperature reflect changes in viscosity.
  Discontinuity with product temperature reflect phase changes in the unfrozen fraction. Exploit in a study of the glass transition and/or eutectic melt of the unfrozen fraction.
Onset and end point of ice crystallization gives rate of ice formation ()
Pre-nucleation data (MW relaxation) predicts the nucleation temperature ()
and (+ soln visc.) control the size distribution of ice crystals and .

## Slide 36: Longinus et al # Poster 18
Mannitol crystallization & melt back
https://doi.org/10.21253/DMU.9759497

## Slide 37: Acknowledgements, Recent Projects & Collaborators
De Montfort University, School of Pharmacy
  Evgeny Polygalov: co-inventor of TVIS instrument
  PhD students : Yowwares Jeeraruangrattana, Bhaskar Pandya, Longinus Ogogua
  Irina Ermolina. Senior Lecturer
National Institute for Biological Standards and Control
  Paul Matejtschuk (IgG TVIS data)
Biopharmaceutical Stability at
Room Temperature
Analytical Technologies for the
Stabilization of Biopharmaceuticals
Government Support for industry
LyoDEA
Lyophilization process analytics
By dielectric analysis
@LyoGroupDMU
Our data
Our Twitter Page
Our WebPage

## Slide 38
Smith, G., Jeeraruangrattana, Y., Ermolina, I. (2018). The application of dual-electrode through vial impedance spectroscopy for the determination of ice interface temperatures, primary drying rate and vial heat transfer coefficient in lyophilization process development. European Journal of Pharmaceutics and Biopharmaceutics, 130 (9), pp. 224-235
Smith, G., Arshad, M.S., Polygalov, E., Ermolina, I., McCoy, T.R., Matejtschuk, P. (2017). Process Understanding in Freeze-Drying Cycle Development: Applications for Through-Vial Impedance Spectroscopy (TVIS) in Mini-pilot Studies. Journal of Pharmaceutical Innovation, 12 (1), pp 26-40
Arshad, M.S., Smith, G., Polygalov, E., Ermolina, I. (2014). Through-vial impedance spectroscopy of critical events during the freezing stage of the lyophilization cycle: The example of the impact of sucrose on the crystallization of mannitol. European Journal of Pharmaceutics and Biopharmaceutics, 87 (3), pp. 598-605

## Slide 39
Smith, G., Arshad, M.S., Polygalov, E., Ermolina, I. (2014). Through-Vial Impedance Spectroscopy of the Mechanisms of Annealing in the Freeze-Drying of Maltodextrin: The Impact of Annealing Hold Time and Temperature on the Primary Drying Rate. Journal of Pharmaceutical Sciences, 103 (6), pp. 1799-1810.
Smith, G., Arshad, M.S., Nazari, K., Polygalov, E., Ermolina, I. ; Taylor, J., Page, T. (2014) Through-Vial Impedance Spectroscopy: A New In-Line Process Analytical Technology for Freeze Drying. Pharmaceutical Technology, 38 (4), pp. 38-46.
Smith, G., Arshad, M.S., Polygalov, E., Irina Ermolina, I. (2014) Factors Affecting the Use of Impedance Spectroscopy in the Characterisation of the Freezing Stage of the Lyophilisation Process: the Impact of Liquid Fill Height in Relation to Electrode Geometry. AAPS PharmSciTech, 15 (2), pp 261–269.
Smith, G., Arshad, M.S., Polygalov, E. and Ermolina, I. (2013) An application for impedance spectroscopy in the characterisation of the glass transition during the lyophilization cycle: The example of a 10% w/v maltodextrin solution. European Journal of Pharmaceutics and Biopharmaceutics, 86 (3 Part B), pp. 1130-1140.
Smith, G., Polygalov, E., Arshad, M.S., Page, T., Taylor, J., Ermolina, I. (2013) An impedance-based process analytical technology for monitoring the lyophilisation process. International Journal of Pharmaceutics, 449 (1-2), pp. 72-83